�������������������������Ŀ

��������������������������ϸ

��    Application File     ��������������������������������������

��                         ���������������������������������������

��

��            Minarad Spectrum's Interactive Manual of

��

��             IR RADIOMETRY and OPTICS Calculations

��

��             By Nelson Engborg and Paul Murphy

��

��            Technical Staff, Minarad Spectrum, Inc.

��

��      Application File  Version 3.12a   08-01-1995

(c) Copyright 1994, 1995 Minarad Spectrum, Inc.

(c) Copyright 1981, 1995 PCSCC, Inc.

All rights reserved.

Read and understand   WARRANTY INFORMATION  . Highlight this line and

type J to skip to this information.

Instructions on using help file are at end of file. Type K to read.

����������͸

� CONTENTS �����������������������������������������������������������

L 30

0.00                  Table of Contents

*    To quickly jump to Sections, position the Section line number

*    on the �JUMP� line (line 12) and type J or type Lxxx (enter)

*    where xxx is line number. For example, to jump to line 139,

*    type L139 (ent).  Type S or R to skip back/forth to Sections.

SECTION                   TITLE                              PAGE�LINE

0.00 ... Table of Contents ................................... 1    31    L 40

1.00 ... Introduction ........................................ 3   117

2.00 ... Acknowledgment of Nelson Engborg, Pioneer in IR ..... 5   197

Radiometry.

3.00 ... About Minarad Spectrum. ............................. 6   232

4.00 ... Minarad Spectrum Infrared Product Line at a ......... 7   311

Glance.

------------------------------------------ 0.00 -- Table of Contents ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------    2 --

5.00 ... Fuzzy Logic Infrared Temperature Spectroscopy ....... 9   385

(FLIRTS)

6.00 ... An Introduction to Calculations Using .............. 10   417

AstroFormCalc 3.                                                 L 50

7.00 ... Methods and Functions Useful in Radiometric ........ 12   529

Calculations.

8.00 ... Electromagnetic Radiation: The IR Source ........... 14   611

(Blackbody)

9.00 ... Calculation of Blackbody Power in a Specified ...... 16   706

Band.

10.0 ... Estimation of Radiance Uncertainty from BB ......... 18   784

Temperature Variation Across Its Radiation

Surface.

11.0 ... Behavior of the Aspherical Refraction Surface ...... 19   840    L 60

(Cartesian Ovoid).

12.0 ... Behavior of the Spherical Refraction Surface ....... 19   865

(Paraxial Rays).

13.0 ... Thin Lenses, Lensmaker's Formula, Gaussian ......... 21   931

Formula and Newtonian Formula.

14.0 ... Thin lenses, Focal Planes and Lateral or ........... 23  1022

Transverse Magnification.

15.0 ... Compound Thin Lenses, Effective Focal Lengths ...... 24  1090

and Dioptric Power.

16.0 ... Thick Lenses, Principal Planes and the Paraxial .... 25  1125    L 70

Approximation.

17.0 ... Thick Lenses, Design of Relay Lens for an IR ....... 26  1151

Spectrometer.

18.0 ... Compound Thick Lenses, Principal Planes and ........ 27  1199

Effective Focal Lengths.

19.0 ... Reflections from Planar, Aspherical and ............ 27  1234

Spherical Mirrors.

20.0 ... Cassegrain and Gregorian Mirrors. .................. 29  1305

21.0 ... Bar Target Dimensions, Cycles/milliradian, ......... 30  1354

Collimator FOV.                                                  L 80

22.0 ... Field of View (FOV) of IR Instrument. .............. 31  1410

23.0 ... Calculating Solid Angles from Fields of View, ...... 33  1506

F-Numbers and Numerical Apertures.

24.0 ... General Signal-to-Noise Calculation for an ......... 34  1528

Radiometer or Photometer at a Specific

Wavelength and Bandwidth.

25.0 ... Signal-to-Noise Calculation of an IR Scene ......... 38  1703

Generator- Radiometer Combination at a Specific

Wavelength and Bandwidth.

26.0 ... Figure of Merit of an IR System: Noise ............. 39  1757    L 90

Equivalent Power Density (NEPD), Radiance (NEN)

and Temperature (NETD).

------------------------------------------ 0.00 -- Table of Contents ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------    3 --

27.0 ... Calculations for Calibrating a Radiometer for ...... 43  1967

Transmission.

28.0 ... In-depth Design of a Spot Radiometer. .............. 46  2098

29.0 ... In-depth Design of a Remote Sensing IR ............. 48  2166

Thermometer.

30.0 ... In-depth Design of a Filter Radiometer for Cloud ... 49  2220

Viewing.

31.0 ... Calculation of Decibel Loss and %Transmittance ..... 55  2516   L 100

of Fibers

32.0 ... Circumference of an Ellipse: Calculation by ........ 56  2547

Series and Integration

33.0 ... Glossary of Terms .................................. 57  2601

34.0 ... MSI Encourages Users to Submit Calculations. ....... 58  2637

35.0 ... Product License and Warranty. ...................... 58  2668

36.0 ... User Support. ...................................... 60  2734

37.0 ... Order/Registration Form ............................ 60  2759

38.0 ... Optics and Infrared References. .................... 62  2825

39.0 ... Instructions for Manually Installing Software ...... 63  2876   L 110

Package.

40.0 ... Help File: Summary of File Viewing Commands ........ 64  2918

���������������͸

� INTRODUCTION  �������������������������������������������������������

1.00          Introduction

������������������������������������������������������������������������

To change the display colors to high-intensity white on blue      L 120

type Q when the highlight bar is lighted on the EQU below.

#

#    HWHITE BLUE

#

This document is an on-line application file which discusses

the design of infrared instrumentation.  Fully-solved problems,

which are developed in an interactive spreadsheet format, form

the theoretical basis for the design of infrared                  L 130

instrumentation including Uv, Vis and IR spectroradiometers,

spectrophotometers, IR sources (blackbodies), target

projectors, collimators, test sets and much more.  The manual

has been featured in "Photonics Spectra" and "Lasers and

Optronics".

Emphasis is directed towards estimation of signal-to-noise as

the primary basis for evaluating the feasibility of the design

----------------------------------------------- 1.00 -- Introduction ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------    4 --

of an opto-electronic instrument.  Excellent references, which

can serve both as a supplement to an existing course in Optics,   L 140

IR and Radiometry or as stand-alone study or self-help guides

are listed at the end of this document in the REFERENCE

section.

Radiometry deals with the measurement of radiation such as UV,

Vis and IR "light", microwaves and (radio frequencies) RF.

Geometrical optics deals with the behavior of "light" (called

Vis and closely related radiations like UV and IR in situations

where the wave nature of the light is negligible.  Radiometry

L 150

The technical staff of Minarad Spectrum has developed this

application to provide guidance to both students and

professionals who work with Infrared (IR) Radiometry and

Optics.  The calculations discussed here can provide help in

mastering the execution of the necessary algebra and math upon

which Optics and UV Vis and IR Radiometry is heavily based.

The staff of Minarad Spectrum will continue to update this file

with new discussions and calculations.  All registered users

will be kept updated about releases of new versions.  The user

is advised to register.  See below.                               L 160

The application is driven by a versatile spreadsheet/calculator

call AstroFormulaCalculator 3 which exists as the file

ACAL3.EXE.  Minarad Spectrum has obtained a distribution

license for ACAL3.EXE and its on-line helpfile ACAL3HLP.HLP.

from PCSCC Inc.

ACAL3.EXE is a compact but very powerful program that will run

on any PC from the XT on up and on any DOS version from 2.1 up.

It contains a highly structured on-line helpfile with countless   L 170

interactive examples on how to setup calculations.  This

program is specifically tailored for analysis based on

algebraic expressions, formulas or equations.  One of its main

functions is to eliminate the "programming-type overhead" in

working with known formulas and to allow meaningful learning to

proceed rapidly.  It is also a virtually paperless program.

Its many and versatile operations can be learned from its

highly structured on-line helpfile with countless interactive

examples on how to setup calculations.

L 180

AstroFormCalc 3 can also be used with any traditional outline

manual or text to develop equation templates that solve

specific problems.  For technical support of AstroFormCalc 3,

the user is required to register for AstroFormCalc 3,

----------------------------------------------- 1.00 -- Introduction ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------    5 --

separately.  See instructions in its on-line helpfile (type key

F1 from the equations display).

This file is organized by topics with are summarized in the

Table of Contents.  The use of AstroFormCalc 3 is discussed in

the on-line help file which is accessed by typing (F1) from the   L 190

equations display.  This entire file can be searched using the

auto-search command "A".

�������������������͸

� NELSON ENGBORG    ����������������������������������������������������

2.00 Acknowledgment of Nelson Engborg, Pioneer in IR Radiometry.

������������������������������������������������������������������������

Special acknowledgment is given to Nelson Engborg coauthor of     L 200

this manual.  His 40+ years of theoretical and practical IR

radiometer experience is visible throughout this manual.

Nelson was probably one of the first researchers to automate

the radiometer design procedures using primitive computer punch

cards!  Over the years, he has enhanced and migrated his design

formulae to and from many calculators and finally to

sophisticated spreadsheets such as 123 and AstroCalculator 3.

One of the first thermal images of a man, which happens to be

Nelson, has been included in this package as the file             L 210

NELSON.PCX.  The image first appeared on the cover of "Applied

Optics" in 1962 and is used with permission of the Optical

Society of America.  At the top of this thermal image is

Nelson's warmed-up '52 Ford.  He had just arrived at Barnes

Engineering in Stamford, CT USA.  At the bottom, is Nelson

smoking his pipe (Note, he doesn't smoke anything now).  Nelson

would like to point out that his ears are darker (lower

temperature) because nicotine is a basal constrictor which

shrinks capillaries!  How's that for some practical medical

spin.  The FLIR camera used a single mirror for two dimensional   L 220

scanning and a pyro-type detector.  The output voltage was used

to modulate a glow tube and the image was ultimately recorded

on a Polaroid film.

If interested, you can see Nelson by viewing the file

NELSON.PCX using the Paintbrush application of MicroSoft's

Windows.

-- 2.00 -- Acknowledgment of Nelson Engborg, Pioneer in IR Radiom... ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------    6 --

�����������������������͸

� MINARAD SPECTRUM INFO ������������������������������������������������   L 230

3.00         About Minarad Spectrum.

������������������������������������������������������������������������

As a matter of background, Minarad Spectrum was founded in 1982

in Fairfield, CT by a group of technology experts which

included Nelson Engborg.  Our company designs and manufactures

a complete line of electro-optical products including:

Radiometers, spectroradiometers, Vis/IR spectrophotometers, IR    L 240

thermometers and IR sources (blackbodies).

Radiometers measure absolute radiance or irradiance

relative to their internal reference blackbody over an

entire wavelength interval.  Spectroradiometers perform

the same functions but at many discrete wavelengths across

a spectral interval.  IR spectrophotometers usually

illuminate the target and measure the transmittance or

reflectance of the target relative to a standard over many

discrete wavelengths.  IR thermometers actually measure      L 250

radiant energy but display an equivalent blackbody

temperature.

IR sources (blackbodies), heat-only and thermoelectric

heat-cool models, conical cavity and extended area radiation

plates, absolute and differential modes,

The IR source or blackbody emits a spectral radiance

distribution which closely approximates that predicted by

Planck's function.  It is the primary reference for most     L 260

IR radiometry research and calibration.

IR test sets, collimators, target wheels, target projectors and

targets/masks,

IR test sets usually consist of an IR source with an

aperture wheel or variable slit, and optionally a filter

wheel, optical modulator (chopper) and possibly a shutter.

A collimator (off-axis parabolic mirror style) is used to

project parallel IR energy with a FOV which depends on the   L 270

largest widths of target masks located in its focal plane.

A target projector consists of an IR test set, target

wheel, collimator and a series of apertures/targets/masks

located in its focal plane.

IR microscopes and telescopes with remote sensing applications,

------------------------------------ 3.00 -- About Minarad Spectrum. ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------    7 --

IR microscopes are used for measuring small spots (10

microns) at relatively short distances (0.5 to 3 inches,

typical).  IR telescopes are used to study larger spots (1   L 280

or 2 feet diameters) at very long distances ( 500 feet

or more).  Radiance, temperature or IR spectroscopy

measurements result.

Hand-portable Vis, NIR and SWIR spectrophotometers,

Compact, hand-portable spectrophotometers and

spectroradiometers for field and lab use.  High

sensitivity and fast scanning.  Spectral Range:  300 to

2500 nm.  Tungsten-halogen source.  Optical Modulator.       L 290

High resolution (5 nm) monochromator.  UV-enhanced Si and

PbS detectors.  Fiber optics accessories.  Comprehensive

Windows Spectroscopy Applications software.  Fully GUI.

Supports extensive calculations including colorimetry,

resolution enhancement, peak fitting, database search and

management, multivariate analysis.

Upgrades of older Barnes Engineering Corp (BEC) radiometers

such as the Mark II 12-550 spectroradiometer.

L 300

Many of the calculations presented in this application manual

are used in the design and operation of these instruments.

If you need any guidance related to these instruments, please

contact the Minarad Staff at (203) 371-6348 or fax a note to

(203) 371-6349. Your support is appreciated.

�����������������������͸

� MINARAD PRODUCT LINE  ������������������������������������������������

L 310

4.00         Minarad Spectrum Infrared Product Line at a Glance.

������������������������������������������������������������������������

IR Sources (Blackbodies).

RBB-20-xxx  Cylindrical conical cavity IR sources

Temperature ranges from 50 to 1000 C and diameters of 1

inch RBB-20-1000 and 2 inches RBB-20-1200.

Emissivities approaching 1. An IR calibration standard.

RBB-30-xxx  Square Radiation plate sources (heat-only).           L 320

Temperatures ranges from 30 to 600 C and dimensions from

2x2 inch to 12 x 12 inch. Textured surfaces for maximum

-------- 4.00 -- Minarad Spectrum Infrared Product Line at a Glance. ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------    8 --

diffusivity.

RBB-40-xxx  Square Radiation plates sources (Thermoelectric

control)

Temperature ranges from -5 to 100 C (lower temperatures

available with enhanced cooling options.) Dimensions from

2x2 to 12x12 inch. Textured surfaces for maximum

diffusivity. Temperature regulation better than 0.01 C.

Very high surface uniformity.                                L 330

RBB-50-xxx Similar to RBB-40-xxx with addition of differential

temperature control. A separate temperature sensor is

attached to a target such that the RBB-50-xxx series PID

controller maintains a constant differential temperature

between the target background and its illumination from

the source. Commonly used in IR target projectors.

Vis-IR Spectroradiometers.

SA-1200-xxx  Cassegrain Spectroradiometer. 4 to 8 inch primary    L 340

mirrors.  2 to 8 mr FOV's.  Multi-segment CVF (circular

variable filter).  Low noise T/E and LN-cooled detectors.

Industry standard PC's.  Complete DOS spectroradiometric

software package.  Complete LabVIEW for Windows (tm,

National Instruments) software package.  Configured as a

spectroradiometer with all interface and calculation vi's.

Develop your own spectroradiometer starting with our

complete set of LabVIEW vi's. Add LabVIEW's powerful

analysis, statistics, display and database management as

you need them! Many models are available. Fully              L 350

calibrated for both small field and large field

measurements using MSI's high performance IR sources.

Contact MSI for more information.

IR Microscope for Microspot Temperature Measurement.

MR-100 Standard metallurgical microscope integrated with a

microspot IR radiometer.  Spot sizes down to 10 microns.

Operating temperature ranges from 30 to 1500 C.  Fully

calibrated.                                                  L 360

Target Projectors, Off-axis Collimators, Apertures, Target Masks

Mirror diameters to 14 inches.  Complete systems including

mirrors, aperture and target wheels, sources and factory

calibration.  Many custom configurations.  Contact Minarad

for more details.

-------- 4.00 -- Minarad Spectrum Infrared Product Line at a Glance. ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------    9 --

Hand-portable Vis-SWIR Spectrophotometer

L 370

VIR-20 Field portable.  Battery operated fast-scanning

Fastie-Ebert monochromator with internal Tungsten Halogen

Quartz illuminator.  Spectral Range 300 to 2500 nm.

Spectral resolution of 5 nm.  Accepts accessories such as

fiber probes and integrating spheres.  Comprehensive

Windows GUI software package including many advanced

spectroscopic functions such as peak fitting, database

creation, search and management and multivariate analysis.

Includes documented database of over 450 Vis-SWIR

signatures.                                                  L 380

����������������������������͸

� FUZZY LOGIC IR TEMPERATURE �������������������������������������������

5.00         Fuzzy Logic Infrared Temperature Spectroscopy (FLIRTS)

������������������������������������������������������������������������

Minarad's commitment to R&D has led to the development of

innovative noncontact temperature measurement technology which

extracts temperatures independent of target                       L 390

emissivity/reflectivity.  MSI has developed a fuzzy logic

approach which has a basis of two very powerful member

functions:  1) a preprocessor function which quickly estimates

multiple surface temperatures and accompanying

emissivities/reflectivities and 2) a statistical member

function which refines the preprocessor estimates.

The FLIRTS technology is applicable to many problems.  For

example, measuring the temperature and emissivity of very small

targets (hot pixels) that are too small to physically attach      L 400

conventional contact sensors.  Also, measuring the temperature

and emissivity of extremely hot targets which are not

compatible to contact sensors such as molten metals.  FLIRTS

allows temperature measurement while surface emissivity varies.

Emissivity measurements independent of a prior knowledge of the

temperature may add a whole new dimension to quality and

reliability measurements in fields such as microelectronic

circuits and to medical diagnostics in the characterization

of biological tissues such as tumors.                             L 410

This technology will be offered on all Minarad radiometers.

----- 5.00 -- Fuzzy Logic Infrared Temperature Spectroscopy (FLIRTS) ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   10 --

�������������������͸

� HOW TO CALCULATE  ����������������������������������������������������

6.00     An Introduction to Calculations Using AstroFormCalc 3.

������������������������������������������������������������������������

This Section provides an introduction to numbers, number          L 420

formats, variables, variable names, mathematical operations,

equations and related topics.  In general, in order to avoid

ambiguities, one must always express numbers, mathematical

expression and equations in an acceptable way when entering

them into a computer program.  These requirements lead to the

formulation of a set of rules which are summarized in the

on-line Help File for the ASTRO 3 program.  This help file can

be reached from the main equation display by typing the (F1)

key.

L 430

For example:  It is very common to have to evaluate an

expression which might be written as:

456.789 x 0.345

As written, one would interpret this expression as:  Find the

product that results from multiplying the decimal 456.789 by

the decimal 0.345.  Because the letter symbols are most often

associated with variable names, the use of "x" as a

mathematical operator for multiplication and the                  L 440

equally-acceptable use of "x" as a variable name in an equation

such as "y=x" would result in confusion for the

expression-interpreter of the program.  As a result, the

multiplication operator is commonly the "*" symbol and the

"by-the-rules" way to enter the above expression would be

456.789 * 0.345 (enter) In the language of "computerese",

"(enter)" means to type the enter key of the keyboard and

usually results in some sort of computer action.

The following examples illustrate some common ways to write       L 450

mathematical expressions and the "by-the-rules" way they are

entered into the ASTRO 3, and many other, computer programs.

These examples exist as equations templates that can be

"called-up" by scrolling the screen until the highlighted bar

appears over the appropriate 'EQU line.

After the templates are loaded, follow the instructions and

always type (F2) from the main equation display to return to

this application helpfile immediately after the 'EQU reference.

L 460

----- 6.00 -- An Introduction to Calculations Using AstroFormCalc 3. ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   11 --

Topic # 1: ENTERING NUMBERS and EXPRESSIONS

Discussion: The following examples show common representations for

expressions and the proper way to enter these into this program. These

expressions can be entered at the ENTER variable below.

commonly used     ASTRO 3 way     Result          Reason

24.1 x 56.7       24.1*56.7      1366.47        * means multipl.

L 470

5

-            5/180          2.77E-2 or    / means division

180                          2.77*10^-2

6 � 3 � 4      6*3*4          72           * not � for multipl.

2(-7)5         2*-7*5         -70            implied multipl.

not allowed.

25�5           25/5           5            / is division operat.   L 480

2      5

14 -  + 8 -  (14+2/7)+(8+5/9)   22.84           express compound

7      9                                     fractions as sums

#                  � ** If highlighted, type Q to get equations that **    L 490

#    IRRAWRIT      � ** solve this problem. Scroll down to highlight.**

Topic # 2: ENTERING EXPONENTS and SQUARE ROOTS

Discussion: The following examples show common representations for

expressions and the proper way to enter these into this program. These

expressions can be entered at the ENTER variable below.

commonly used     ASTRO 3 way     Result          Reason               L 500

___

�144         (space) sqr(144)  12         use sqr(..) function

to take square root.

2.3

10             10^2.3           199.5       ^ means raise to power

4

----- 6.00 -- An Introduction to Calculations Using AstroFormCalc 3. ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   12 --

5.67 x 10      5.67e4            56700      e means base 10

4     3

3x10 -2x10     3e4-2e3 or        28000      calculation in

3*10^4-2*10^3                scientific notation.     L 510

2 -1

(2x10 )        (2e2)^-1          0.005    no implied exponentiation

___

� 6

------    (space) sqr(6)/(4*sqr(3))   .3535  must match parentheses

4�3

L 520

#                  � ** If highlighted, type Q to get equations that **

#    IRRAEXPO      � ** solve this problem. Scroll down to highlight.**

�������������������͸

� USEFUL FUNCTIONS  ����������������������������������������������������

7.00     Methods and Functions Useful in Radiometric Calculations.

������������������������������������������������������������������������   L 530

Nelson Engborg considers two mathematical techniques very

useful for general radiometric analysis.  These techniques are

1) Integration and 2) Solving complex equations for independent

variables by numerical methods.  AstroCalculator 3 can be used

to integrate one or more equations.  AstroCalculator 3

integrates active equations which have their dependent

variables followed by an equal sign "=".  Show below are three

equations which may be entered into the program using the F6

key command.  Following the equation entry, the main equation     L 540

screen should look similar to what is shown below.

Y1=X                              Y1=0

Y2=2*X                            Y2=0

Y3=4*X                            Y3=0

X = 0

Note: equations Y1, Y2 and Y3 are active. Typing G to

integrate with a range from 0 to 2 will yield this solution:      L 550

Y1=X                              Y1=2

-- 7.00 -- Methods and Functions Useful in Radiometric Calculatio... ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   13 --

Y2=2*X                            Y2=4

Y3=4*X                            Y3=8

X = 0

To direct integration to only Y2, move the cursor to Y1 and

type I.  Next, move the cursor to Y3 and type I.  Finally, type

G to integrate with the same range from 0 to 2.  The results      L 560

are now:

Y1=X                              Y1 0

Y2=2*X                            Y2=4

Y3=4*X                            Y3 0

X = 0

Note: only the equation Y2, which is followed by the equal

sign, is active and integrated.                                   L 570

To load these equations and practice, type Q when the EQU line

is highlighted.

#

Y1=X Y2=2*X Y3=4*X

#    Y1=0 Y2=0 Y3=0 X=0

#

Type (F2) to return to application file.

L 580

AstroCalculator 3 can be used to numerically backsolve.

Assume that the equation is a polynomial such as

Y1=3-4*X+2*X^2                                  Y1=0

X=0

To find the value of X at which Y1 equals 9, numerically

backsolve this equation using the S command.  Move the cursor

to X and type S.  For the desired value of Y, type 9 (enter).     L 590

Use the default initial search range, just type (enter).  The

program will solve for Y1=9 at X=-1.  Another solution may

exist. Set X=30 and repeat. X=3 is also found!

To load these equations and practice, type Q when the EQU line

is highlighted.

#

-- 7.00 -- Methods and Functions Useful in Radiometric Calculatio... ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   14 --

Y1=3-4*X+2*X^2

#    Y1=0  X=0                                                             L 600

#

Type (F2) to return to application file.

A complete discussion of these functions, with interactive

examples, can be found in the AstroCalculator 3 on-line help

file which can be read by typing the (F1) key.

������������͸

� RADIATION  ����������������������������������������������������������

L 610

8.00       Electromagnetic Radiation: The IR Source (Blackbody)

������������������������������������������������������������������������

Electromagnetic radiation propagates through space in

wavefronts of fluctuating magnetic and electric fields.  The

frequency, or wavelength, of this radiation identifies it as

radio waves, microwaves, infrared, visible UV and so on.  The

first equation template identifies the kind of radiation from

its wavelength or frequency.  Geometrical optics and Radiometry

is most applicable to radiation in and around the visible (Vis)   L 620

which includes ultra-violet (UV) and infrared (IR) radiations.

Topic # 3: WAVELENGTH, RADIATION TYPE, FREQUENCY

Problem # 3.1: Enter wavelength in microns.

(a) What radiation has a wavelength of 400 nm (400E-3 microns)?

(b) What radiation has a frequency of 400 THz (400*10^14 Hz)?

#                  � ** If highlighted, type Q to get equations that **

#    IRRARAIN      � ** solve this problem. Scroll down to highlight.**    L 630

An IR source (blackbody) is the "perfect" absorber of all

incident electromagnetic radiation.  It is also the "perfect"

radiator because it is in equilibrium at a fixed temperature

and so must emit all its incident energy.  However, the

wavelength distribution of its emitted radiation follows

Planck's Law, Setfan's Law and Wein's Law.  Emittance is mainly

in the infrared region and is felt as heat, hence the term

"blackbody".  Above 600�C objects begin to glow "red" as the      L 640

emitted energy moves into the visible region.  Understanding

these Laws is fundamental to understanding radiation in

general.  However, the distribution of illuminance from a

visible source rarely shows the IR-source distribution.  For

------- 8.00 -- Electromagnetic Radiation: The IR Source (Blackbody) ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   15 --

example, a tungsten filament heated to 3000 K will show a

"graybody" Planckian spectral distribution, but only in the

transmission region of its quartz enclosure.

Topic # 4: RADIATION, POWER EMITTED vs ABSORBED, STEFAN'S LAW, BLACKBODY

L 650

Problem # 4.1: A household-type 'radiator' has emissivity of 0.6 and

area of 2.0 m^2

(a) At what rate is radiation absorbed by the radiator when the walls

of the room are at 27�C?

(b) At what rate is radiation emitted from the radiator when its

temperature is 50�C?

(c) What is the net radiation rate from radiator?

#                  � ** If highlighted, type Q to get equations that **

#    IRRA1946      � ** solve this problem. Scroll down to highlight.**    L 660

The relative radiance or power in any wavelength interval can

be calculated by integrating Planck's equation.  The total

power is found from Stefan's law which is the closed integral

of Planck's equation over all wavelengths.  The following

calculation shows how one estimates the percentage of blackbody

radiation in any wavelength relative to the maximum

theoretically possible.

L 670

Topic # 5: PLANCK'S LAW, EMISSIVE POWER, RADIATION IN WAVELENGTH

INTERVAL

Discussion: Calculate the emissive power in the far IR (1000 to 5.

6microns) middle IR (5.6 to 1.5 micron) and near IR (1.5 to 0.72

microns) for objects at 1000 and 100 �C.

#                  � ** If highlighted, type Q to get equations that **

#    IRRAPLLW      � ** solve this problem. Scroll down to highlight.**

L 680

Finally, one can estimate the radiation cone of an IR source

which is its "field of view".  Its FOV is based on a

point-source located at the end of its cavity.  The rays are

limited, or stopped, by its exit aperture.  One half of its FOV

is its divergence angle.  IR will irradiate anything within the

cone.  If a plane surface is placed perpendicular to the

optical axis at a given working distance from the source, it

will cast a spot on that plane.  Note, atmospheric effects are

not included.                                                     L 690

------- 8.00 -- Electromagnetic Radiation: The IR Source (Blackbody) ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   16 --

Topic # 6: SPOT DIAMETER, SIZE, FOCUS AT INFINITY, CONE ANGLE, WORKING

DISTANCE, IR POINT SOURCE

Discussion: What is the spot size ( = spot diameter) at a working

distance of 2 inches from a cylindrical IR source with an aperture

diameter of 1 inch and a cone angle of 12�?

#                  � ** If highlighted, type Q to get equations that **

#    IRRASFOV      � ** solve this problem. Scroll down to highlight.**    L 700

������������͸

� IR SOURCE  ����������������������������������������������������������

9.00       Calculation of Blackbody Power in a Specified Band.

������������������������������������������������������������������������

The following example is included to illustrate a common

calculation used to estimate the temperature and diameter of a    L 710

BB source (cylindrical-conical cavity). MSI's sales

representative requested the MSI staff to propose an IR source.

A customer required a blackbody (BB) IR source that would

provide 0.5 W over the 3-5 um band (a popular IR short wave

band for thermal imaging cameras also sometimes called FLIR's).

MSI reviewed his request for the BB source.  The MSI staff made

the following calculations in order to determine the

appropriate BB temperature to yield the 0.5 W emitted energy.

MSI used both 1 and 2 inch exit apertures which correspond to     L 720

our RBB-20-1000 and RBB-20-2000 cylindrical conical cavity

sources.  These IR cavity sources are very good approximations

to the ideal BB with emissivities greater than 0.98.

The results of MSI's calculations are summarized here.  The

formula template used to calculate these values is presented

below.

T�C     Exit H          Lambda       % of       Power   Power

irradiance      Maximum      SB         W       W           L 730

W/cm�           (um)         max H      1 inch  2 inch

(3-5um)                                 apert.  apert.

500     0.6696          3.74         33         3.4     13.6

400     0.3296          4.30         28         1.7     6.7

-------- 9.00 -- Calculation of Blackbody Power in a Specified Band. ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   17 --

300     0.129           5.05         21         0.65    2.6

----

200     0.0352          6.12         12         0.18    0.71        L 740

----

MSI would recommend our RBB-20-1000 at T=300�C or our

RBB-20-2000 at 200�C.  Both will deliver approximately 0.5 W

(See Table above).  The larger source RBB-20-200 offers a

greater radiation surface of approximately 20 cm� vs 5 cm� for

the smaller RBB-20-1000 source.

The calculation template used to estimate the emitted energy in

the 3-5 um spectral band as a function of BB temperature is

included below.  Planck's equation is integrated to yield the     L 750

spectral irradiance.  Follow the instructions included with the

template.

#                  � ** If highlighted, type Q to get equations      **

#    IRRASELB.EQU  � ** discussed above.  Scroll down to highlight.  **

Other useful calculations of energy in specific bands follow:

L 760

Topic # 7: PLANCK'S LAW, EMISSIVE POWER, RADIATION RATIO VS TEMPERATURE

Discussion: Calculate the ratio of emitted power in the 8 - 12 thermal

IR band for perfect blackbody radiation at temperatures of 25 and

100�C.

#                  � ** If highlighted, type Q to get equations that **

#    IRRATVSW      � ** solve this problem. Scroll down to highlight.**

L 770

Topic # 8: PLANCK'S LAW, EMISSIVE POWER, RADIATION RATIO, WAVELENGTH

BAND

Discussion: Calculate the ratio of emitted power in the 4 - 6 micron

band to that in the 8 - 12 micron band at a temperature of 1000�C.

#                  � ** If highlighted, type Q to get equations that **

#    IRRARVSW      � ** solve this problem. Scroll down to highlight.**

L 780

-------- 9.00 -- Calculation of Blackbody Power in a Specified Band. ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   18 --

������������͸

� IR SOURCE  ����������������������������������������������������������

10.0       Estimation of Radiance Uncertainty from BB Temperature

Variation Across Its Radiation Surface.

������������������������������������������������������������������������

Minarad currently offers large area radiation sources with 12 x

12 inch plates with two heat/cool control options:  the

RBB-30-120 series (heat only) from 50 to 325�C and the            L 790

RBB-40-120 (thermoelectric controlled) from approximately -10

to 100�C.  The RBB-30-120 uses an integral PID algorithm to

control the heat-only rods while the RBB-40-120 uses the PID to

control heal/cool thermoelectric modules.  The RBB-30-120 is a

cost-effective solution to generate high temperatures over

large areas.  However, the RBB-40-120, with its T/E modules and

its capability to add/remove heat, tightly regulates

temperature and maintains much higher temperature uniformity

across the radiation plate.  For example, MSI's thermal imaging

studies suggest that the T/E RBB-40-120 will maintain better      L 800

than 0.1�C uniformity across most of its surface from -10 to

100�C while the heat-only RBB-30-120 will maintain �2�C

uniformity across a 5x5 inch square at its center at 300�C!

Recently a MSI customer requested MSI to estimate the effect of

uniformity on the IR source radiance for the T/E RBB-30-120 at

the heat-only RBB-30-120 at 100�C.

Wave    Radiation       Active       Temp       Radiance

(mic)   Plate ance      Area         Variation  Variation       L 810

Size            in x in      �C          %

in x in

1-2     12 x 12         5 x 5        �2         4.9

3-5     12 x 12         5 x 5        �2         2.19

8-12    12 x 12         5 x 5        �2         0.99

The equation template used to make these calculations is          L 820

presented below for your review.  After loading the template,

follow instructions on the procedure for integration by typing

the comma key to read the comment after loading the template.

Topic # 9: IR SOURCE RADIANCE VS TEMPERATURE UNIFORMITY

Problem # 9.1:

(a) How will a change of 2�C at an operating temperature of 300 �C in

----- 10.0 -- Estimation of Radiance Uncertainty from BB Temperature ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   19 --

the 1-2 micron spectral band affect the total radiance?

(b) In the 3-5 micron band?                                               L 830

(c) In the 8-12 micron band?

#                  � ** If highlighted, type Q to get equations that **

#    IRRAUNIF      � ** solve this problem. Scroll down to highlight.**

������������������������͸

� ASPHERICAL REFRACTION  �����������������������������������������������

11.0       Behavior of the Aspherical Refraction Surface (Cartesian        L 840

Ovoid).

������������������������������������������������������������������������

The following problem studies the behavior of a spherical wave

diverging from a point source and impinging on an aspherical

refracting surface known as a Cartesian Ovoid.  These surfaces

include the ellipsoid, hyperboloid and paraboloid.

Topic # 10: ASPHERICAL REFRACTING SURFACES, CARTESIAN OVOID

L 850

Problem # 10.1: Write an expression for the Cartesian ovoid with image

and object distances of 15 and 12 cm, respectively. Ovoid made of

glass Ng=1.5 and surrounded by air Na=1.

(a) Calculate the ray length from ovoid surface to its image if that

point is 14 cm from the source object.

(b) What does a negative distance mean?

#                  � ** If highlighted, type Q to get equations that **

#    IRRA402       � ** solve this problem. Scroll down to highlight.**

L 860

������������������������͸

� SPHERICAL REFRACTION   �����������������������������������������������

12.0       Behavior of the Spherical Refraction Surface (Paraxial

Rays).

������������������������������������������������������������������������

The basis of most geometrical optics is the spherical

refracting surface which is of higher symmetry than its           L 870

aspherical cousins, discussed above.  This higher symmetry

allows many basic formulas to be derived.  The following

problems deal with spherical surfaces.

----- 12.0 -- Behavior of the Spherical Refraction Surface (Paraxial ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   20 --

Topic # 11: SPHERICAL REFRACTING SURFACES, PARAXIAL RAYS

Problem # 11.1: The input end of a wide glass (Ng=1.6) fiber is ground

to a convex hemisphere of 1.5 cm radius.

(a) If a point source is located 4 cm to the left of the hemisphere's

vertex, where will its image be?                                          L 880

#                  � ** If highlighted, type Q to get equations that **

#    IRRA405       � ** solve this problem. Scroll down to highlight.**

Topic # 12: SPHERICAL REFRACTING SURFACES, PARAXIAL RAYS

Problem # 12.1: The input end of a wide glass (Ng=1.6) fiber with

convex hemisphere of 1.5 cm radius is dipped in water (Nw=1.32).

(a) If a point source is located 4 cm to the left of the vertex, where    L 890

will its image be?

#                  � ** If highlighted, type Q to get equations that **

#    IRRA406       � ** solve this problem. Scroll down to highlight.**

Topic # 13: SPHERICAL REFRACTING SURFACES, OBJECT FOCAL LENGTH

Problem # 13.1: An interface separates air on the left and glass (Ng=1.

44) on the right.                                                         L 900

(a) What is the object focal length if its radius of curvature is 6

cm?

#                  � ** If highlighted, type Q to get equations that **

#    IRRA408       � ** solve this problem. Scroll down to highlight.**

Topic # 14: SPHERICAL REFRACTING SURFACES, IMAGE FOCAL LENGTH

Problem # 14.1: An interface separates air on the left and glass (Ng=1.    L 910

44) on the right.

(a) What is the image focal length if its radius of curvature is 6 cm?

#                  � ** If highlighted, type Q to get equations that **

#    IRRA408A      � ** solve this problem. Scroll down to highlight.**

Topic # 15: SPHERICAL REFRACTING SURFACES, OBJECT FOCAL LENGTH

Discussion: What is the radius of curvature of a glass fiber (Ng=1.48)     L 920

----- 12.0 -- Behavior of the Spherical Refraction Surface (Paraxial ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   21 --

required to focus light rays 95 cm in front of its spherical interface

in a non-aqueous media of isopropanol (Nm=1.38)?

#                  � ** If highlighted, type Q to get equations that **

#    IRRA409       � ** solve this problem. Scroll down to highlight.**

�������������͸

� THIN LENSES ����������������������������������������������������������

L 930

13.0       Thin Lenses, Lensmaker's Formula, Gaussian Formula and

Newtonian Formula.

������������������������������������������������������������������������

The majority of calculations in geometrical optics make two

basic assumptions:  1) the impinged-upon surface is a sphere

(spherical approximation) and 2) the impinging light strikes

that surface at small angles with respect to its optical axis

(paraxial approximation).  In general practice, however, one

usually invokes these formulas for any lens calculation.  The     L 940

following problems deal with "thin" lenses.  Such a lens is

infinitely thin and all distances are measured from their

interior principal plane which passes through its center.

An illustration of a "thin" lens is shown below.  Its center is

marked by "||".  All distances are measured from it.

ss

s  s

s  s                     M[T] + means erect image      L 950

� Y[O]           s    s

�������o���������s�||�s��������o����������o  S source point

S     F[O]       s    s       F[I]        �  P image point

s  s               Y[I] �  S[O] Source distance

|�X[O]�|���f����� s  s�����f���|����X[I]��|  S[I] Image distance

ss

|--------S[O]------|-----------S[I]-------|

Sign Convention:  X[O] + means its left of F[O]        focal_length < 0

X[I] + means its right of F[I]           means           L 960

Y[I], Y[O] + means above optical axis        diverging lens.

The following problems illustrate a number of calculations

involving thin lenses.

Topic # 16: THIN SPHERICAL LENS EQUATION, LENSMAKER'S FORMULA

----- 13.0 -- Thin Lenses, Lensmaker's Formula, Gaussian Formula and ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   22 --

Discussion: The source and image points are located 5 and 7 cm,

respectively, from a spherical concave lens with source radius of -3

cm and image radius of 4 cm. What should the index of refraction of       L 970

the glass be?

#                  � ** If highlighted, type Q to get equations that **

#    IRRA410       � ** solve this problem. Scroll down to highlight.**

Topic # 17: THIN SPHERICAL LENS EQUATION, LENSMAKER'S FORMULA

Problem # 17.1: A source is 20 cm from a glass (ng=1.6) convex-plano

lens with a radius of curvature of 4 cm. Where is the image located       L 980

(a) when its flat surface faces source and

(b) when its flat surface is away from source?

#                  � ** If highlighted, type Q to get equations that **

#    IRRA411       � ** solve this problem. Scroll down to highlight.**

Topic # 18: THIN SPHERICAL LENS EQUATION, FOCAL LENGTH vs N[G], N[L]

Discussion: Calculate the focal length of a bi-concave thin lens made      L 990

of glass (Ng=1.6) and immersed in isopropanol (Na=1.35). The source

radius is -9 and the image radius is 18 cm.

#                  � ** If highlighted, type Q to get equations that **

#    IRRA413       � ** solve this problem. Scroll down to highlight.**

Topic # 19: LENS EQUATION, FOCAL LENGTH vs INDEX REFRACTION, RADIUS CURV

Problem # 19.1: A bi-concave (R1=-8 cm, R2=12 cm) thin lens is made of    L 1000

hepa-fluor (Ng=1.37) and is immersed in solvent alpha-k-1 (Nm=1.77).

(a) What is its focal length?

(b) is it converging or diverging?

#                  � ** If highlighted, type Q to get equations that **

#    IRRA414       � ** solve this problem. Scroll down to highlight.**

Topic # 20: MEASURING INDEX REFRACTION in MEDIUM from FOCAL LENGTHS

L 1010

Problem # 20.1: A bi-convex lens of index 1.6 has a focal length of 60

cm in air and 240 cm when immersed in a transparent fluid.

----- 13.0 -- Thin Lenses, Lensmaker's Formula, Gaussian Formula and ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   23 --

(a) What is the index of refraction of the transparent fluid?

#                  � ** If highlighted, type Q to get equations that **

#    IRRA416       � ** solve this problem. Scroll down to highlight.**

�������������������͸

� THIN LENS IMAGING ����������������������������������������������������  L 1020

14.0       Thin lenses, Focal Planes and Lateral or Transverse

Magnification.

������������������������������������������������������������������������

Predicting or estimating the location of an object, also called

a source or target, and the location of its image are very

important.  The following problems deal with this topic.

Topic # 21: THIN LENS IMAGERY, CONVEX, TRANSVERSE MAGNIFICATION           L 1030

Problem # 21.1: A thin bi-convex lens of index 1.6 has equal radii of

40 cm in air.

(a) Compute the focal length.

(b) If the image is to be projected full-size on a white wall, what

are the image and object distances?

#                  � ** If highlighted, type Q to get equations that **

#    IRRA419       � ** solve this problem. Scroll down to highlight.**

L 1040

Topic # 22: THIN LENS IMAGERY, CHARACTERISTICS for REAL OBJECTS,

GAUSSIAN

Problem # 22.1: A 10 cm tall statue of Curly Howard is sitting 100 cm

from a positive (converging) lens of 30 cm focal length.

(a) Describe Curly's image.

#                  � ** If highlighted, type Q to get equations that **

#    IRRA420       � ** solve this problem. Scroll down to highlight.**   L 1050

Topic # 23: THIN LENS IMAGERY, CHARACTERISTICS for REAL OBJECTS,

NEWTONIAN

Problem # 23.1: A 9 cm tall statue of Mo Howard is sitting 100 cm from

a positive (converging) lens of 30 cm focal length.

(a) Describe Mo's image.

-------- 14.0 -- Thin lenses, Focal Planes and Lateral or Transverse ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   24 --

#                  � ** If highlighted, type Q to get equations that **   L 1060

#    IRRA421       � ** solve this problem. Scroll down to highlight.**

Topic # 24: THIN LENS IMAGERY, MAGNIFICATION, OBJECT, IMAGE DISTANCES

Problem # 24.1: An object positioned 38 inches to the right of a

positive lens is imaged 8 inches to its left.

(a) Where will the object appear if its image is moved to 2 inches

from the lens?

(b) Describe it.                                                         L 1070

#                  � ** If highlighted, type Q to get equations that **

#    IRRA423       � ** solve this problem. Scroll down to highlight.**

Topic # 25: THIN LENS IMAGERY, FIND F from R1, R2, BI-CONVEX, GAUSSIAN

Problem # 25.1: A thin bi-convex glass (N[g]=1.6) with radii of 20 and

70 cm casts a half-sized image of a plant pot on a white wall.

(a) What is its focal length?                                            L 1080

(b) What are the lens to plant and lens to wall distances?

#                  � ** If highlighted, type Q to get equations that **

#    IRRA425       � ** solve this problem. Scroll down to highlight.**

����������������������͸

� COMPOUND THIN LENSES �������������������������������������������������

15.0       Compound Thin Lenses, Effective Focal Lengths and              L 1090

Dioptric Power.

������������������������������������������������������������������������

Lens combinations are very useful.  The following problems deal

with combinations of "thin" lenses.

Topic # 26: COMPOUND THIN LENSES, MAGNIFICATION, FRONT BACK FOCAL

LENGTHS

Problem # 26.1: A compound lens consists of two bi-convex lenses of       L 1100

focal lengths 15 and 30 cm, respectively, separated by 60 cm.

(a) Describe the image generated by a statue 10 cm tall placed 12 cm

in front of the first lens.

(b) What separation distance is required for unity magnification.

---------- 15.0 -- Compound Thin Lenses, Effective Focal Lengths and ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   25 --

#                  � ** If highlighted, type Q to get equations that **

#    IRRA427       � ** solve this problem. Scroll down to highlight.**

Topic # 27: COMPOUND THIN LENSES, MAGNIFICATION, MICROSCOPE               L 1110

Problem # 27.1: A low cost microscope is fabricated out of a bottom

lens of 2 cm focal length, a 10 cm cardboard tube and a top lens with

a 8 cm focal length.

(a) Locate the image of an ant 3 cm from the bottom lens.

(b) Describe the ant.

#                  � ** If highlighted, type Q to get equations that **

#    IRRA431       � ** solve this problem. Scroll down to highlight.**

L 1120

��������������͸

� THICK LENSES ���������������������������������������������������������

16.0 Thick Lenses, Principal Planes and the Paraxial Approximation.

������������������������������������������������������������������������

The basic thin lens formalism and equations can be applied to

"thick" lenses if effective focal lengths and principal planes

(from which distances are measured) are properly defined.  The   L 1130

single principal plane of the thin lens is split into two

principal planes H1 and H2 separated by the lens thickness D.

The next problem deal with this topic.

Topic # 28: THICK LENS, GAUSSIAN FORMULA, PARAXIAL APPROXIMATION

Problem # 28.1: A double convex lens has radii of 6 and 24 cm, a

thickness of 3 cm and index of 1.6.

(a) Locate both the principal and focal points and compute the image

distance for an object 15 cm in front of V1.                             L 1140

(b) determine the magnification of the image and

(c) determine the f.f.l. and b.f.l.

#                  � ** If highlighted, type Q to get equations that **

#    IRRA442       � ** solve this problem. Scroll down to highlight.**

L 1150

-- 17.0 -- Thick Lenses, Design of Relay Lens for an IR Spectrome... ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   26 --

��������������͸

� THICK LENSES ���������������������������������������������������������

17.0 Thick Lenses, Design of Relay Lens for an IR Spectrometer.

������������������������������������������������������������������������

The MSI VIR-20 Vis-SWIR (short-wave Infrared) is a

comprehensive and versatile hand-portable

spectrometer/spectroradiometer which provides a spectral range

from 300 to 2500 nm.  Spectral selection employs Fastie-Ebert

grating monochromators.  The external collectors are fiber

light-guide bundles with circular apertures at remote-sensing

probe interfaces and slit-like apertures at the spectrometer     L 1160

interfaces.  This arrangement allows virtually any sensor

(transmission cell, integrating sphere, immersion probe,

reflectance probe, etc.) to be interfaced to the VIR-20.  The

collected energy (as an approximate 4x0.6mm slit-like image) is

focused directly on the entrance slit of the visible

monochromator and is folded, via a steering mirror, to the

entrance slit of the SWIR monochromator.

For the SWIR channel the image must be transferred over an

additional 60 mm path length.  A bi-convex lens is used to       L 1170

transfer the energy via 1-to-1 image transfer optics.  The next

calculation shows how to calculate the proper radius of

curvature for this lens as a function of the lens center

thickness, index of refraction and total path length.  The

actual lens is chosen with an effective focal ratio (f-number)

lower or "faster" than the system f-number of approximately

F/3.  For example, an F/1 lens would be 15 mm in diameter if

the efl (effective focal length) of the thick lens is about 15

mm, as it is in the next example.  As always, the calculated

radius of curvature and optical layout must allow for some       L 1180

adjustment in order to be compatible with the lens

manufacturer's tooling, assembly and alignment and so on.

Topic # 29: THICK LENS, IR RELAY, BI-CONVEX, IMAGE TRANSFER

Problem # 29.1: A double convex lens is used as a relay lens for image

transfer across a 60 mm path. CaF2 stock with index of 1.426 at 1.5

microns is chosen.

(a) What is the radius of curvature?

(b) What is its effective focal length?                                  L 1190

#                  � ** If highlighted, type Q to get equations that **

#    IRRARLAY      � ** solve this problem. Scroll down to highlight.**

-- 17.0 -- Thick Lenses, Design of Relay Lens for an IR Spectrome... ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   27 --

�����������������������͸

� COMPOUND THICK LENSES ������������������������������������������������

18.0       Compound Thick Lenses, Principal Planes and

Effective Focal Lengths.                                         L 1200

������������������������������������������������������������������������

As with "thin" lenses, "thick" lenses can also be used in

combinations.  It should be noted that if the thickness of the

individual lenses, as measured between their principal planes,

approaches 0.0, then the solutions are the same as for the

"thin" lens combinations discussed above.  The following

problems deals with this topic.

Topic # 30: LENS COMBINATIONS, THICK LENSES, GAUSSIAN FORMULA             L 1210

Discussion: Two identical bi-convex lenses are separated by 30 mm. Each

lens has a radius of 50 and 30 mm, a thickness of 20 mm and an index

of 1.6. Calculate individual and system focal lengths.

#                  � ** If highlighted, type Q to get equations that **

#    IRRA443       � ** solve this problem. Scroll down to highlight.**

Topic # 31: LENS COMBINATIONS, THICK LENSES, GAUSSIAN FORMULA             L 1220

Discussion: A bi-concave lens (radius=-50 mm, index=1.6) is positioned

150 mm up a plastic tube in front of a plano-convex lens (radius 50 mm,

index 1.6). Calculate the image location of a 9 mm carpenter ant 90 mm

in front of the device.

#                  � ** If highlighted, type Q to get equations that **

#    IRRA446       � ** solve this problem. Scroll down to highlight.**

L 1230

�������������͸

� MIRRORS     ���������������������������������������������������������

19.0       Reflections from Planar, Aspherical and Spherical

Mirrors.

������������������������������������������������������������������������

Lenses are refracting surfaces and mirrors are reflecting

surfaces.  There exists a correspondence between the equations

that predict the behavior of mirrors and those of lenses if one  L 1240

takes into account sign conventions, etc.  A spherical convex

mirror is shown below.  Mirror surface is represented by 's'.

---------- 19.0 -- Reflections from Planar, Aspherical and Spherical ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   28 --

The convex mirror has a "virtual" focal point and image.  The

virtual image lies behind or "in" the mirror between the mirror

and the focal point.  The concave mirror has both the image and

object in the "real" domain on the same side of the mirror.  In

general, image/object distance are considered positive if they

are "real" and negative if they are virtual.  Thus, the convex

mirror has one positive and one negative distance while the

concave mirror has two positive distances.                       L 1250

�

�      s

� o�                   S = source object

�  �    s �                   V = vertex of sphere

�  �          s   � P     C          P = image (between F and C)

S ��� � � � � � � � �o� � ��o � � o          R = 2 * F

|                   V|s     |  ��

|                    | s   �|�  R      S[O] = distance S to vertex

|                    |  s�  |         S[I] = distance vertex to image     L 1260

|-------S[O]---------|-S[I]-|    C = center of sphere R = radius curvat.

Sign Convention: S[O] + means its left of vertex V,

S[I] + means its left of vertex V,

R    + means center C is right of V

Convex mirror has negative focal length

Concave mirror has positive focal length

The following problems deal with mirrors.

L 1270

Topic # 32: SPHERICAL REFRACTING MIRRORS, PARAXIAL RAYS

Problem # 32.1: A convex mirror (17 cm radius) casts the image of a

bird on a screen 20 cm distant.

(a) Where is the bird located and

(b) what will it look like?

#                  � ** If highlighted, type Q to get equations that **

#    IRRA460       � ** solve this problem. Scroll down to highlight.**

L 1280

Topic # 33: SPHERICAL REFRACTING MIRRORS, IMAGE vs MAGNIFICATION

Problem # 33.1: A spherical mirror casts an inverted half-sized image

of an object that is located 110 in from its vertex.

(a) Where will the image be located?

#                  � ** If highlighted, type Q to get equations that **

---------- 19.0 -- Reflections from Planar, Aspherical and Spherical ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   29 --

#    IRRA461       � ** solve this problem. Scroll down to highlight.**

L 1290

Topic # 34: SPHERICAL REFRACTING MIRRORS, IMAGE vs MAGNIFICATION

Problem # 34.1: An eraser 5 cm tall sits 150 cm in front of a concave

mirror of focal length 400 cm.

(a) Describe the image.

#                  � ** If highlighted, type Q to get equations that **

#    IRRA495       � ** solve this problem. Scroll down to highlight.**

L 1300

������������͸

� CASSEGRAIN �����������������������������������������������������������

20.0       Cassegrain and Gregorian Mirrors.

������������������������������������������������������������������������

Mirror optics offer broadband spectral characteristics and are

useful as collectors and as transfers in instruments designed

for broadband operation such as the Minarad SA-200 and SA-1200   L 1310

spectroradiometers.  The Cassegrain and Gregorian are very

similar.  They both have concave primary mirrors.  The

Cassegrain has a diverging or convex secondary while the

Gregorian has a converging or concave secondary.  The

Cassegrain style with its diverging secondary mirror allows a

longer optical path from the detector to the secondary mirror.

Thus, it is more accommodating to the radiometer's IR hardware

such as Minarad's inline blackbody reference cavity, field stop

and filter/CVF.  Minarad uses the Cassegrain in its SA-200 and

SA-1200 spectroradiometers                                       L 1320

The following problems deal with Cassegrain and Gregorian

mirror systems.

Topic # 35: CASSEGRAIN OPTICS, SPHERICAL THIN MIRRORS, MAGNIFICATION

Discussion: A Cassegrain optic has a large concave mirror of FL=12 in.

and smaller convex of FL=4 in., respectively, initially separated by

10 in. Focus the system on a 3 inch tall target 42 inches in front of

the primary and describe the image located 2 inches behind the primary   L 1330

mirror's hole.

#                  � ** If highlighted, type Q to get equations that **

#    IRRACASS      � ** solve this problem. Scroll down to highlight.**

-------------------------- 20.0 -- Cassegrain and Gregorian Mirrors. ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   30 --

Topic # 36: GREGORIAN OPTICS, SPHERICAL THIN MIRRORS, MAGNIFICATION

Problem # 36.1: A Gregorian optic has a large concave mirror of FL=12

in. and smaller convex of FL=6 in., respectively, initially separated    L 1340

by 8 in.

(a) Focus the system on a 3 inch tall target 42 inches in front of the

primary and describe the image located 3 inches behind the primary

mirror's hole.

#                  � ** If highlighted, type Q to get equations that **

#    IRRAGREG      � ** solve this problem. Scroll down to highlight.**

L 1350

���������������͸

� FIELD OF VIEW ��������������������������������������������������������

21.0 Bar Target Dimensions, Cycles/milliradian, Collimator FOV.

������������������������������������������������������������������������

MSI specializes in the design of target projectors (TP).  TP's

find use in calibration of thermal imaging equipment and in MTF

(Modulation Transfer Function) instruments.  TP's are

electro-optical devices which consist of a collimator (usually   L 1360

an off-axis parabolic mirror type with exit apertures ranging

from 5 to 14 inches in diameter).  Aperture and pattern wheels

position various masks at the collimator's focal point.  Masks

are illuminated by an IR source (blackbody) which projects the

mask's pattern from the collimator to the unit under test

(UUT).  MSI supplies complete TP systems including collimator,

aperture/target wheels, IR sources and optics table.  For

example, the MSI TP-5-30 system would have a collimator with a

5 inch off axis parabolic mirror and 2x2 inch square targets.

L 1370

Aperture targets are usually illuminated with circular BB's

such as the MSI RBB-20-1000 1 inch cylindrical-conical source

with operating temperatures from 50 to 1000�C.  2x2 or 4x4 inch

masks are more commonly illuminated with square radiation-plate

sources such as the MSI RBB-30 series (heat only) or the RBB-40

series (thermoelectric heat/cool varieties).

Very often one must calculate the dimensions of the bar

targets, also called Nyquist targets, used to measure spatial

properties of optics such as the minimum detectable and          L 1380

-- 21.0 -- Bar Target Dimensions, Cycles/milliradian, Collimator ... ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   31 --

resolvable temperature difference (MDTD and MRTD).  It should

be noted that a "bar target" with one bar is more commonly

called a "slit" target if it has a high aspect ratio or called

a uniform-illumination target if it has an aspect ratio near

unity.  The spatial frequency of bar patterns is also closely

related to the spacing of pixels in focal plane array (FPA)

detectors.

The calculation, included below, allows one to calculate the

total dimensions of each bar target (usually with 4 bars) as a   L 1390

function of the spatial frequency in cycles/milliradian.  In

addition, the minimum FOV (field of view) at the collimator's

exit aperture is also calculated.  If the actual FOV is less

than this value, the projected pattern will be clipped.

Topic # 37: TARGET DIMENSIONS, TARGET PROJECTOR, COLLIMATOR, FOCAL

LENGTH, SPATIAL FREQUENCY, MILLIRADIANS, N BAR PATTERN (N=4, TYPICAL)

Discussion: Calculate dimensions of a 4-bar Nyquist target with 7/1

aspect ratio and freq of 0.53 cycles/milliradian. Targets are to be      L 1400

placed at the focal point of a 70 inch collimator.

#                  � ** If highlighted, type Q to get equations that **

#    IRRABARC      � ** solve this problem. Scroll down to highlight.**

���������������͸

� FIELD OF VIEW ��������������������������������������������������������

22.0       Field of View (FOV) of IR Instrument.                          L 1410

������������������������������������������������������������������������

What an IR instrument "sees" is usually characterized by its

field of view, FOV.  The FOV is twice its acceptance angle.  If

an IR detector is located at the focal point of an optical

system containing a single converging (convex lens), the

illustration, below, can be used to identify terminology:

FOV/2 is 1/2 of FIELD OF VIEW  L 1420

� �   <----- FOV/2                    or ACCEPTANCE ANGLE

�......�.....�

S �            |    �     �          |

P �    OPTIC-->|              �      |<-- DETECTOR

O �            |          �   �      |

T �...... .... �    �                     Detector at focal point of

---------------------- 22.0 -- Field of View (FOV) of IR Instrument. ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   32 --

� �    �                                optical system.

|---WORK_DIS-|--------S[DET]-------|

|---------FL----------|

L 1430

Detector:  The IR detector's diameter, if circular, or the

length of its sides, if square or rectangular.

S[DET]:  Distance from the center of the lens to the detector.

OPTIC:   Diameter of the lens  FL:  The focal length of lens

WORK_DIS:  the distance from the center of the lens to a plane,

perpendicular to the optical axis and containing an image of

the detector called its "spot".

SPOT: The diameter of the image of the detector at WORK_DIS.

FOV: Field of view in degrees, Acceptance angle is FOV/2

L 1440

The shape of the spot is assumed to be the same as that of the

detector.  A circular detector has a circular spot (and the

same FOV in all directions perpendicular to the optical axis).

A rectangular one has a rectangular spot (and two FOV's), a

square detector has a square spot (but only one FOV).

If the detector is a perfect point, then the systems FOV is 0�.

For any sized detector, a collimating location exists, measured

from the lens, at S[DET] = FL + F# * DIAM_DET.  For example, a

1 mm detector in an F/2 system will have a 0� FOV if the         L 1450

detector is located 2 mm from the focal point.  A negative FOV

implies that the detector's image is focussed in front of the

lens.

The following equations can be used to calculate FOV's working

distances, spot sizes and F# for any simple system with a

detector at an arbitrary location behind the lens.

Topic # 38: FOV OPTICAL SYSTEM, DETECTOR VS DISTANCE FROM LENS

L 1460

Problem # 38.1:

(a) Calculate the FOV of a radiometer with a 0.3 inch lens, a 0.6 inch

FL and a 4 mm detector 8 mm behind its focal point.

(b) What is the spot diameter at 400 inches in front of the lens?

#                  � ** If highlighted, type Q to get equations that **

#    IRRAFDET      � ** solve this problem. Scroll down to highlight.**

The next set of equations can be used to make system             L 1470

calculations when the detector is located at the optical focal

point, a very common location in IR systems.

---------------------- 22.0 -- Field of View (FOV) of IR Instrument. ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   33 --

Topic # 39: FOV OPTICAL SYSTEM, DETECTOR AT FOCAL POINT

Problem # 39.1:

(a) Calculate the FOV of a radiometer with a 0.3 inch F/2 lens and a 4

mm detector at its focal point.

(b) What is the spot diameter at 4 inches in front of the lens?

L 1480

#                  � ** If highlighted, type Q to get equations that **

#    IRRAFFOC      � ** solve this problem. Scroll down to highlight.**

The FOV (or acceptance or divergence angle) and lens diameter

can be measured experimentally by measuring the spot diameter

vs.  the working distance and using linear regression.  The

following example illustrates this procedure.

Topic # 40: SPOT DIAMETER vs WORKING DISTANCE, FOV, DIAMETER OPTIC,       L 1490

LINEAR REGRESSION

Discussion: A quality department wants to measure its laser's

divergence angle (FOV/2) and its exit aperture by measuring the spot

diameter vs working distance and using linear regression. If data are

(11 mm, 10 ft), (23,20), (49, 50) and (95.1, 100), what are the FOV

and aperture diameter?

#                  � ** If highlighted, type Q to get equations that **

#    IRRASVSW      � ** solve this problem. Scroll down to highlight.**   L 1500

�������������͸

� SOLID ANGLE ����������������������������������������������������������

23.0 Calculating Solid Angles from Fields of View, F-Numbers and

Numerical Apertures.

������������������������������������������������������������������������

The following equation template can be used to calculate the     L 1510

field of view, f-number and numerical aperture for a given

solid angle in steradians an also to calculate the solid angle

given the field of view, f-number or numerical aperture.

Topic # 41: SOLID ANGLE CALCULATOR, F#, NUMERICAL APERTURE, FOV

Discussion: What is the solid angle viewed by an IR detector with a FOV

of 17.45 mr? To solve, type S, then (end esc) FOV_MR (ent), then (end

-- 23.0 -- Calculating Solid Angles from Fields of View, F-Number... ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   34 --

esc) 17.45 (ent). Use default search range, type (ent).

L 1520

#                  � ** If highlighted, type Q to get equations that **

#    IRRASANG      � ** solve this problem. Scroll down to highlight.**

�����������������͸

� S/N RADIOMETER  ������������������������������������������������������

24.0       General Signal-to-Noise Calculation for an Radiometer or

Photometer at a Specific Wavelength and Bandwidth.

������������������������������������������������������������������������  L 1530

The design of a radiometer or photometer usually follows from

fundamental calculations which attempt to "predict" the

signal-to-noise ratio (S/N) over a specified spectral

bandwidth.  The "fundamental" calculations can be separated

into at least two "sub-fundamental" calculations:  1)

Estimation of the radiant flux (power) incident on the

radiometer collecting head and 2) Estimation of the power that

is incident of the radiometer detector cell.

L 1540

The following calculations are modeled on a Minarad SA-1200

spectroradiometer with a 4 inch Cassegrain collecting optic, 4

mr half-power field of view (FOV), various transfer optics,

monochromator (5 nm bandwidth) and a lead sulfide (PbS)

detector cell(1-2.5 micron range).  The measurement wavelength

is 1.5 microns and the optical modulator (chopper) runs at 400

Hz. The electronic bandwidth is 100 Hz.

MSI supplies similar SA-1200 spectroradiometer for use as

transfer radiometer in many metrological applications.           L 1550

Configurations vary. The use of a circular variable filter

(CVF) with 2% spectral bandwidth and a LN-cooled InSb detector

(1-5 micron range) is common.

Step 1: Estimation of Photonic Energy at Radiometer Collector

Head.

The source is assumed to be a MSI RBB-20-1000 1 inch diameter

cylindrical-conical cavity IR source masked with a small (0.094

inch diameter) aperture.  This blackbody is assumed to           L 1560

approximate a diffuse emitter (Lambertian) point source which

propagates a constant radiance over its 180� total divergence

angle.  The radiated energy of the source is a function of its

temperature.  For visible wavelength calculations, use the

-- 24.0 -- General Signal-to-Noise Calculation for an Radiometer ... ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   35 --

correlated color temperature.  The following template is used

to calculate the flux (power) incident on the radiometer.

Topic # 42: POWER AT RADIOMETER COLLECTOR, SMALL FIELD

Discussion: Calculate the power in Watts across a 4 inch collector of a   L 1570

radiometer with 4 mr half-power FOV (field of view) at 1.5 micron (5

nm BW). The source is a diffuse emitter at 800�C ( diameter 0.094 inch

at 8 feet).

#                  � ** If highlighted, type Q to get equations that **

#    IRRASRC1      � ** solve this problem. Scroll down to highlight.**

The collected energy is transferred to the detector by 4 mirror

bounces (2 in the Cassegrain and 2 in the transfer optics),      L 1580

through the monochromator and through the detector window.

These elements introduce an optical attenuation factor which is

called the optical efficiency.  Its proper specification is

critical to the S/N.  However, it value tends to be somewhat

subjective.

There are two sources of noise:  Photonic noise, which results

from the motion of electrons or current flow in the detector

and Detector noise which is specified by the manufacturer as D*

(Dee star) or NEP (noise-equivalent power).  Detectors are also  L 1590

characterized by an electrical resistance and by a responsivity

which scales photonic flux to electrical voltage or current.

The following template is used to calculate the signal-to-noise

above background (small field approximation) at the detector

for a specified amount of collected power. The detector output

must be offset or nulled so that the radiometer "sees 0

voltage" when looking at an ambient-temperature target.

Topic # 43: SIGNAL TO NOISE AT DETECTOR                                   L 1600

Discussion: Calculate the S/N (signal to noise) of a PbS detector at

the focal point of a radiometer with a 0.45 thruput. Detector

performance specs are: Size: 1x3 mm, Electrical bandwidth: 100 Hz,

D_star:2.57e-11 Jones, Resistance: 1.45 MOhm, Responsivity: 1.19e6 V/W.

Assume 5.11e-9 Watts are collected by a 4 inch Minarad SA-1200

Cassegrain Radiometer.

#                  � ** If highlighted, type Q to get equations that **

#    IRRADET1      � ** solve this problem. Scroll down to highlight.**   L 1610

-- 24.0 -- General Signal-to-Noise Calculation for an Radiometer ... ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   36 --

The next template combines the calculations of the source

energy and detector S/N so that the S/N can be estimated as a

function of the source temperature and aperture size.

Topic # 44: COMPLETE SIGNAL TO NOISE CALCULATION: SOURCE TO DETECTOR

Discussion: Calculate the S/N at the PbS detector for a 4 inch

Cassegrain radiometer with 4 mr half-power FOV (field of view) at 1.5    L 1620

micron (5 nm BW). The source is a diffuse emitter at 800�C( diameter 0.

094 inch at 8 feet).

#                  � ** If highlighted, type Q to get equations that **

#    IRRATOT1      � ** solve this problem. Scroll down to highlight.**

To fill the radiometer field of view, change (increase,

usually) the source aperture size until the solid angle from

the source equals the radiometer FOV solid angle.  Filled field  L 1630

of view measurement at higher temperatures can easily saturate

the radiometer.

The next template will calculate full field of view

signal-to-noise estimates.

Topic # 45: COMPLETE SIGNAL TO NOISE CALCULATION: FULL FIELD

Discussion: Calculate the S/N at the PbS detector for a 4 inch

Cassegrain radiometer with 4 mr half-power FOV (field of view) at 1.5    L 1640

micron (5 nm BW). The source is diffuse at 800�C and fills the

radiometer field of view.

#                  � ** If highlighted, type Q to get equations that **

#    IRRATOT2      � ** solve this problem. Scroll down to highlight.**

For collimated energy, the source is no longer diffuse.  Its

energy (relative to a diffuse emitter with a small aperture) is

essentially compressed into a narrow solid angle and             L 1650

beamed in the direction of the radiometer.  The collimator exit

diameter should be larger than that of the radiometer

collector so that the filled field of view approximation is

valid.  A small aperture is located at the collimator focal

point.  The diameter of this aperture throttles the amount of

energy.  Usually, it must be very small to prevent the

-- 24.0 -- General Signal-to-Noise Calculation for an Radiometer ... ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   37 --

saturation of the radiometer.

The next template calculates S/N for a collimated source.

L 1660

Topic # 46: COMPLETE SIGNAL TO NOISE CALCULATION: COLLIMATED SOURCE

Discussion: Calculate the S/N at the PbS detector for a 4 inch

Cassegrain radiometer with 4 mr half-power FOV (field of view) at 1.5

micron (5 nm BW). The collimated source is at 800�C, aperture size: 0.

094, focal length: 25 in.

#                  � ** If highlighted, type Q to get equations that **

#    IRRATOTC      � ** solve this problem. Scroll down to highlight.**

L 1670

Summary of Signal-to-noise measurements for various sources.

Source to radiometer distance=96 inches, Source temperature

is 800�C. Aperture is 0.094 inches, as required.

Source              Power                 Calculated

at radiometer         Signal-to-noise        L 1680

head (W)

diffuse             5.11e-9               170/1

point

collimated          9.10e-7               27,805/1

source

filled              9.27e-6               184,149/1              L 1690

FOV

Caution:  It is important to realize that the present

calculations address the conversion of the photonic energy to

useful electrical energy at the detector.  Many fundamentally

sound designs have fallen on hard time because of the failure

to properly amplify and digitize the healthy detector signal.

------ 25.0 -- Signal-to-Noise Calculation of an IR Scene Generator- ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   38 --

��������������������͸                                                    L 1700

� S/N SCENE GENERATOR���������������������������������������������������

25.0       Signal-to-Noise Calculation of an IR Scene Generator-

Radiometer Combination at a Specific Wavelength and Bandwidth.

������������������������������������������������������������������������

An IR Scene Generator is used to generate and project moving IR

images.  An IR Target Projector has passive masks and/or

apertures which are placed in the focal plane of a collimator

and are back-illuminated by an IR source (blackbody).  An IR     L 1710

Scene Generator is a specialized version of the basic target

projector that has the passive masks and the source replaced by

a single two-dimensional array of hot pixels.  Pixel dimensions

are typically 50 microns square and upwards of 16384 (128x128

pixels).  Pixels are turned on/off by increasing/decreasing

their temperatures.  Frames rates of 100 Hz are possible.  The

IR radiation is collimated and projected from the instrument.

The transmitted energy can be detected by a thermal imager

(FLIR) or radiometer.

L 1720

Scene Generator          Minarad SA-1200 Cassegrain Radiometer

.    )  ..................

.        )                   )

array     .           )                   )

|-|   .               )                   )

|-|.                  )               ) .           |

|-|   .               )               )        .    | detector

.           )               ) .           |

.        )                   )                         L 1730

.     )                   )

)  .................)

|------ fl ----------|

The following calculation is used to estimate the

signal-to-noise vs.  the number of hot pixels when the energy

is received by a radiometer with a diameter equal to that of

the exit aperture of the scene generator and located directly

at the exit aperture.  The radiometer FOV is assume to be        L 1740

completely filled.

Topic # 47: IR SCENE GENERATOR, SIGNAL-TO_NOISE CALCULATION

Discussion: Calculate the S/N at the MCT detector for a 3.2 inch

Cassegrain radiometer with 1x1 mm detector at 12 microns (2% CVF BW).

Source is collimated IR scene generator with an array of 50 micron hot

pixels (326�C) at focal point.

------ 25.0 -- Signal-to-Noise Calculation of an IR Scene Generator- ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   39 --

#                  � ** If highlighted, type Q to get equations that **   L 1750

#    IRRASCNE      � ** solve this problem. Scroll down to highlight.**

�����������������͸

� FIGURE OF MERIT ������������������������������������������������������

26.0       Figure of Merit of an IR System: Noise Equivalent Power

Density (NEPD), Radiance (NEN) and Temperature (NETD).

������������������������������������������������������������������������

L 1760

The performance of an IR instrument depends on many factors.

Some of the more important are:  field of view (FOV), usually

from 1 mrad to 20�, spectral bandwidth typically from 1.3 to 2

% of its nominal wavelength if a CVF is used, detector type,

such as InSb, MCT or a sandwich combination of these, and

electrical bandwidth (from 1 to 20 KHz or wider, depending on

the acquisition rate).  Performance is usually evaluated in two

different ways depending upon whether the target fills the

detector field of view or whether the target is essentially a

point source within the detector field.                          L 1770

At least three useful figures of merit are available for

evaluating an IR instrument's ability to sense target radiation

above its electrical noise.  1) The noise equivalent power

density (NEPD) which characterizes its ability to detect a

target (essentially a point-source) that does not fill the

detector field of view.  2) The noise equivalent radiance (NEN)

which indicates the instrument's ability to detect a target

that fills the instrument's FOV.  3) The noise equivalent

temperature (NET) which is the smallest different in target      L 1780

temperature that can be detected.

Definitions of calculated and supplied parameters follow:

PEAK_WAVE - The wavelength of maximum spectral radiance as

predicted by the Wein Displacement Law.

FOV[MRAD] - The field of view (FOV) in milliradians.  Usually

the FOV is express in degrees if FOV>0.5� and in mrad if         L 1790

FOV<0.5�.

W_R[STER] - The solid angle subtended by the FOV.  The FOV is

the apex angle of a plane bisecting the FOV cone.  For a FOV

-- 26.0 -- Figure of Merit of an IR System: Noise Equivalent Powe... ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   40 --

that can be approximated as a right-circular cone, the solid

angle is equal to 2*pi*(1-cos(FOV/2)) or pi*tan(FOV/2)^2.

NEP[W] - The noise equivalent power in Watts at the detector.

NEP is calculated from the detector D* (dee-star) at the

specified wavelength, the detector area DT_AREA in mm and the    L 1800

detector bandwidth DT_BW in Hz.

NEP= sqr(DT_AREA*DT_BW)/D_STAR.

NEP[W] is used to supply "the noise" in most figure of merit

calculations.

NEPD[W] - The noise equivalent power density at the entrance of

the collecting optic in Watts�cm�.  The NEPD (also sometimes

called the noise equivalent irradiance or NEI) is the target     L 1810

energy density required to equal the detector noise scaled to

its energy density at the entrance of the collecting optic.

With NEPD, the instrument achieves a signal-to-noise ratio of

1.  The smaller the NEPD, the better the performance of the IR

instrument.  NEPD is derived from a point source target which

does not fill the instruments field of view and is always

measured above the background irradiance.

NEN[W] - The noise equivalent radiance at the entrance of the

collecting optic in Watts�cm��ster.  The NEN is the target       L 1820

radiance required to equal the detector noise scaled to its

equivalent radiance at the entrance of the collecting optic.

With NEN, the instrument achieves a signal-to-noise ratio of 1.

The smaller the NEN, the better the performance of the IR

instrument.  NEN is always derived from a target that fills the

instrument FOV.

N_EFF - the effective radiance at the entrance to the

collecting optics for a blackbody target at temperature TEMP[C]

and in the specified spectral bandwidth SP_BW[HZ]                L 1830

WET[W] - The power at the detector equivalent to N-EFF.

S_N - the system signal-to-noise ratio.  Calculated for a

blackbody target filling its FOV.  It is calculated from the

ratio of the target power at the detector WET to the noise

equivalent power NEP.

NETD[C] - The blackbody temperature difference (always at a

specific temperature TEMP[C]) equal to the detector/electronic   L 1840

-- 26.0 -- Figure of Merit of an IR System: Noise Equivalent Powe... ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   41 --

system noise.  NETD is directly related to the system's NEN and

the filter region/spectral bandwidth of the filter used in the

system.  NETD = NEN / (dNEN/dT at TEMP[C]).

The following parameters are known for the instrument's

intended design.

CL_A[CM�] - The net collecting area in cm� of the entrance

optic reduced for any obscuration such as from the secondary     L 1850

mirror of the Cassegrain collector.

DT_A[MM�] - The area of the detector in square mm.

DT_BW[HZ] - The electrical bandwidth of the system in Hz.

D-STAR - The detector detectivity in cm sqr(Hz)/Watt.  Usually

supplied with the detector specifications.

FOV[�] - The instruments FOV in degrees.  Generally, based on    L 1860

the limiting aperture which may be the detector or a field

stop.

OPTEFF - The optical efficiency of the instrument.  The product

of the individual efficiencies of collecting optics, filters,

transfer optics and demodulating circuits.

SP_BW[U] - The spectral bandwidth in microns(U).  (Usually the

width of the 50% transmission or power points relative to the

center of the filter.)                                           L 1870

TEMP[C] - The target temperature in �C.  The target is always

assumed to be a blackbody, a perfect emitter.

WVLNG[u] - The spectral wavelength in microns.  The center of

the filter band.

An equation template that calculates NEPD, NEN, S_N and NET for

a blackbody target that fills the instrument FOV follows.

L 1880

#                  � ** If highlighted, type Q to get equations      **

#    IRRASENS.EQU  � ** discussed above.  Scroll down to highlight.  **

For an instrument with an InSb detector at 77K and CVF spectral

-- 26.0 -- Figure of Merit of an IR System: Noise Equivalent Powe... ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   42 --

selector with 2% nominal resolution, use this template:

Topic # 48: INFINITY FOCUS, FULL FOV, FIGURE OF MERIT, INSB DET(77K),

2% CVF                                                                   L 1890

Discussion: Calculate NEPD, NEN and NET for an InSb, CVF radiometer

with the following: Detector area 1 sq mm, det bandwidth 1 Hz, optical

efficiency of 0.06, collecting area of 120 sq cm, FOV of 2.5 mr,

target temperature of 90�C at wavelengths of 1, 3 and 5 microns.

#                  � ** If highlighted, type Q to get equations that **

#    IRRAINSB      � ** solve this problem. Scroll down to highlight.**

L 1900

For an instrument with an MCT (HCT, Mercury Cadmium Telluride,

mercad) detector at 77K and CVF spectral selector with 2%

nominal resolution, use this template:

Topic # 49: INFINITY FOCUS, FULL FOV, FIGURE OF MERIT, MCT DET(77K), 2%

CVF

Discussion: Calculate NEPD, NEN and NET for an MCT, CVF radiometer with

the following: Detector area 1 sq mm, det bandwidth 1 Hz, optical

efficiency of 0.06, collecting area of 120 sq cm, FOV of 2.5 mr,         L 1910

target temperature of 90�C at wavelengths of 6, 12 and 14 microns.

#                  � ** If highlighted, type Q to get equations that **

#    IRRAMCT       � ** solve this problem. Scroll down to highlight.**

A plot of estimated signal-to-noise vs wavelength for a Minarad

Spectrum SA-1200 narrow FOV (2.5 mr), CVF radiometer with a 6

inch Cassegrain collector and MCT detector is shown below.

L 1920

1658-                                                        k

|         Signal to noise vs. wavelength              k

|         MSI SA-1200 Radiometer                   k     k

|         6 inch Cassegrain Collector          k           k

|         MCT (77K) detector                k              k

1243-         2.5 mr FOV                      k

|                                       k                    k

|                                     k                      k       L 1930

|                                   k

|                                  k                           k

-- 26.0 -- Figure of Merit of an IR System: Noise Equivalent Powe... ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   43 --

829.-                                k                             k

|                              k                                 k

|                             k

|                            k                                   k

|                          k

414.-                        k

|                       k

|                     k                                              L 1940

|                   k

|                k

+0.00-k k   k   k  k

||--------------|--------------|--------------|--------------|-----

1             4.0            7.0            10.0           13.0

wavelength (microns)

To plot this spectral response use the following template.

Type comma first to read the comment for instructions.           L 1950

Topic # 50: INFINITY FOCUS, FULL FOV, FIGURE OF MERIT, MCT DET(77K), 2%

CVF, PLOT SIGNAL to NOISE VS WAVELENGTH

Discussion: Plot the signal-to-noise for an MCT, CVF radiometer with

the following: Detector area 1 sq mm, det bandwidth 1 Hz, optical

efficiency of 0.06, collecting area of 120 sq cm, FOV of 2.5 mr,

target temperature of 90�C from 1 to 14.9 microns.

#                  � ** If highlighted, type Q to get equations that **   L 1960

#    IRRAPSN       � ** solve this problem. Scroll down to highlight.**

������������������������͸

� RADIOMETER TRANSMISSION�����������������������������������������������

27.0 Calculations for Calibrating a Radiometer for Transmission.

������������������������������������������������������������������������

All values are nominal values.  All equation values are:         L 1970

wavelength (lambda), microns or u; distances, in, mm, cm and m;

and Neff (effective radiance in the pass band); watts/

cm�-ster.  Where a spectral selecting CVF (continuously

variable filter) is defined, the bandwidth is approximately 2%

of instantaneous Lambda peak response.

When atmospheric transmission measurements are desired over a

known measured distance, it has been established that a

-- 27.0 -- Calculations for Calibrating a Radiometer for Transmis... ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   44 --

radiometer and matching size off-axis collimator when placed in

contact, and in perfect alignment, allow generation of a zero    L 1980

absorption Lambda-wavelength transmission file.  This is only

true when the value of irradiance H from the collimator in

contact with the radiometer is the same as the value received

from the collimator placed at the desired operating distance.

It is very useful to simulate long path distances in the

laboratory by stopping down the collimator exit aperture such

that the radiometer can be placed directly in front of the

collimator.  In order to properly calibrate the radiometer, one

must adjust the collimator exit irradiance H by substituting an  L 1990

accurately measured aperture of the correct size to generate

the correct zero-distance value of H.

Formula: H = Neff * A0/(FL+D)^2 = Neff * A1/FL^2

H is the irradiance.  A0 is the exit area of the collimator and

radiometer and D is the desired distance.  A1 is the area of

the aperture after stopping and FL is the collimator focal

length.  The effective target radiance, Neff, is constant at

both zero and desired distances.  Note :  FL << D approximation  L 2000

made.

The following template calculates the stopping diameter of the

aperture.  Note: A Minarad Spectrum C-6-32 collimator has a 6

inch exit aperture and a 32 inch focal length.  For energy

measurements, it typically includes a MSI RBB-20-1000 1 inch

(50 to 1000�C) blackbody as its illuminant.

Topic # 51: TRANSMISSION, COLLIMATOR, SOURCE, APERTURE-EQUIVALENT

IRRADIANCE                                                               L 2010

Discussion: What diameter must the exit aperture of a MSI C-6-32

collimator be stopped down to in order for its irradiance H at zero

distance to be identical to its irradiance at 4 kilometers (without

stopping) as collected by a identical-diameter radiometer (neglect

atmospheric absorbtion effects)?

#                  � ** If highlighted, type Q to get equations that **

#    IRRADIAH      � ** solve this problem. Scroll down to highlight.**

L 2020

The next problem illustrates how the figure of merit of the

collecting radiometer can be estimated when it is at zero

distance in the laboratory rather that at 4 kilometers in the

-- 27.0 -- Calculations for Calibrating a Radiometer for Transmis... ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   45 --

field.  These calculations can be used to test the practicality

of attempting to collect energy from the collimator at 4

kilometers.  The S/N ratio should be 100 or greater if the

experiment is to be practical.

Assume:                                                          L 2030

A Minarad Spectrum C-6-32 6 inch Collimator with an RBB-20-1000

IR source with 1 inch clear opening operating at 1000�C ( an

effective radiance value of 4.390E-2), Lambda peak at 5.0

microns.  CVF bandwidth 2% of peak bandwidth).  Collimator

irradiance H from previous example is 5.003e-11 (See equation

template below).

A Minarad Spectrum SA-1200 6 inch diameter spectral radiometer

with the following specifications:                               L 2040

Collecting area A0 in cm�: 120

FOV (field of view): 5 milliradians

Filter CVF spectral range: 1.8-5.5 microns

Detector 1x1 =1 mm� InSb D* (dee star): 2.0E11

Electrical bandwidth: 10 Hz

System Thruput Analysis:

(Mirrors 0.95 reflectance) ^4 * (CVF transmission 0.4) *

(Synchronous demodulation factor 0.315) *                        L 2050

(Detector lens transmission 0.7)

--yield-->  0.0718 system transmission.

Using the 6 inch MSI C-6-32 collimator previously described we

find the following:

(N.E.P.) = (Det area cm� * bw in Hz)^1/2

-------------------------      = 1.58E-12

D*

L 2060

(N.E.P.D.) =         N.E.P.

---------------------   =   1.835E-13

A0 * System thruput

S./N. Ratio =  H value at the radiometer    5.003E-11

--------------------------  = --------- = 272.6

N.E.P.D.              1.835E-13

The above problem assumes zero transmission losses at 5.0 mic.   L 2070

-- 27.0 -- Calculations for Calibrating a Radiometer for Transmis... ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   46 --

In reality transmission losses occur even in the IR

transmission bands depending upon concentration of absorbing

materials.  These calculations strongly suggest that if the

radiometer is positioned next to the collimator with the proper

stopped-down aperture (See previous example.), the radiometer

should be able to "see" the target in the lab and also in the

field when the radiometer is positioned 4 kilometers away from

the collimator without the stop.

The template which contains a generalized form of this           L 2080

calculation appears next.

Topic # 52: RADIOMETER, SIGNAL-TO_NOISE, COLLIMATOR, SOURCE, IRRADIANCE,

RADIANCE, NEP, NEPD

Discussion: What signal-to-noise is expected for a Minarad Spectrum

SA-1200 6 inch diameter spectroradiometer with an InSb detector and

CVF spectral selector collecting energy at 5 microns from a 6 inch

collimator source at a distance of 4 kilometers?

L 2090

#                  � ** If highlighted, type Q to get equations that **

#    IRRASNEP      � ** solve this problem. Scroll down to highlight.**

�����������������������͸

� SPOT RADIOMETER DESIGN������������������������������������������������

28.0       In-depth Design of a Spot Radiometer.

������������������������������������������������������������������������

L 2100

Most radiometers are focussed at infinity.  Some applications

required the radiometer to be focussed up close.  For example,

We might require a 8 mm spot at working distance of 200 mm.

The spectral region is lambda peak 7.9 microns, bandwidth 0.5

micron with minimum target temperature of 200 �C.  Electrical

bandwidth is 10 Hz.  A Minarad Spectrum spot radiometer or

remote-sensing thermometer has the following specifications:

Optic Collecting Diameter= 1.0 inch (25.4 mm).  Spot size:

(Note:  the spot size is defined as the half-power FOV * 3 or,   L 2110

in this case, 8.0 mm/3) = 2.666 mm FOV size.  Detector:  pyro

1x1 mm, D* (dee star)= 2E8.  effective target radiance Neff:

lambda 7.9, bw 0.5 at 200 �C = 4.2111E-3

Defining lens FL (focal length) as

---------------------- 28.0 -- In-depth Design of a Spot Radiometer. ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   47 --

FOV / obj distance   =  det size / img distance

Solving this proportion equation for img distance yields 75 mm.

The lens focal length:  1/f = 1/obj + 1/img = 55mm .             L 2120

In practice we would use a 50 mm standard lens (F/2) and

readjust the detector to lens distance for best focus at 200 mm

target distance.

The power appearing at the detector is included in the solid

angle At/d^2 with respect to the FOV diameter.  Area of the

target (At)= 5.58E-2 cm�/ distance�= 400 cm� yields target

solid angle of 1.395E-4 ster.  The radiometer lens collecting

area (Ac) is 5.064 cm� Power at the detector is                  L 2130

P = Ac * solid angle * Neff.

Power = 5.064 * 1.395E-4 * 4.211E-3  = 2.9748E-6

System thruput efficiency = lens transmission * filter

transmission * synchronous demodulator factor.  Efficiency =

0.9 * .8 * .315 = 0.2268

(N.E.P.) = (area of detector in cm� * bw in Hz)^1/2              L 2140

---------------------------------       = 6.972E-9

( detector D_star * thruput efficiency)

S/N = Power/N.E.P.  = 426, a very reasonable signal to noise

ratio.

The template which contains a generalized form of this

calculation appears next.

Topic # 53: SPOT RADIOMETER, FOV, DETECTOR SIZE, OBJECT, IMAGE, SIGNAL-   L 2150

TO_NOISE, NEP

Problem # 53.1:

(a) What focal length (1 inch dia. biconvex lens) and detector image

distance are required to produce an 8 mm spot at 200 mm.

(b) What signal- to-noise is expected for a 200�C target using a pyro

detector at 7.9 mics, spectral BW 0.5 mics and electrical BW of 10 Hz?

#                  � ** If highlighted, type Q to get equations that **

#    IRRASPRD      � ** solve this problem. Scroll down to highlight.**   L 2160

---------------------- 28.0 -- In-depth Design of a Spot Radiometer. ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   48 --

����������������������͸

� IR THERMOMETER DESIGN�������������������������������������������������

29.0       In-depth Design of a Remote Sensing IR Thermometer.

������������������������������������������������������������������������

A radiometer is to be calibrated as a remote-sensing

thermometer.  We require a full field of view of 1.146 degrees   L 2170

(20 mrad).  Spectral band of lambda 1.8 - 3.0 microns (peak

lambda=2.4 mic, BW=1.2 mic).  Electrical bandwidth 20 Hz.

Minimum detectable temperature 60�C, Neff (1.8 - 3.0) =

2.70E-6.

The Minarad Spectrum industrial temperature measuring

instrument such as RST-20 remote sensing thermometer has the

following specifications.

Lens (Water free Quartz), 1.0 inch diameter, FL 50 mm.  Filter   L 2180

coated Germanium LP.  Germanium cuts on at 1.8 mic.  and the

detector cutoff is approximately 3.0 microns.  Detector is 1.0

x 1.0 mm single stage cooled PbS, 20% bias.  Power on a PbS or

PbSe detector is approximately 1mw/mm-sq Electrical BW = 20 Hz.

Power at the detector = Ao * instrument solid angle * Neff

Power = 5.06 * 2*PI*SIN(20E-3/2)� * 8.7107E-7 = 8.6E-9

System thruput = lens transmission * filter * synchronous        L 2190

modulator.  Efficiency = .9 * .85 * .315 = 0.2409

(N.E.P.)= (area of detector cm * bw (electrical))^1/2

--------------------------------------   = 4.64E-11

( D_star    * thruput efficiency )

S_N = P / N.E.P. = 185. This value suggests 60�C is near the

minimum useful temperature (S_N=100) of this instrument.

The template which contains a generalized form of this           L 2200

calculation appears next.  The example in this template

actually shows how one can estimate the minimum detectable

temperature of an IR thermometer if one knows something about

how it was built.

Topic # 54: REMOTE SENSING THERMOMETER/ RADIOMETER, FOV, DETECTOR SIZE,

SIGNAL-TO_NOISE, MINIMUM DETECTABLE TEMP

-------- 29.0 -- In-depth Design of a Remote Sensing IR Thermometer. ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   49 --

Discussion: What is the minimum detectable temperature for a PbS IR

thermometer with a 1.146� FOV and 1.8 to 3.0 microns spectral range, 1   L 2210

inch collector and 20 Hz electrical BW?

#                  � ** If highlighted, type Q to get equations that **

#    IRRARST1      � ** solve this problem. Scroll down to highlight.**

���������������������������͸

� ADVANCED RADIOMETER DESIGN��������������������������������������������

30.0 In-depth Design of a Filter Radiometer for Cloud Viewing.            L 2220

������������������������������������������������������������������������

Minarad Spectrum's (MSI) is pleased to submit informal

specifications for a spectral radiometer for temperature

measurements of small areas such as clouds in cold regions of

the sky.

The MSI staff has thoroughly investigated the possibility of

using uncooled detectors, such as pyroelectric.  We have also

looked into the use of our standard 6-inch SA-1000 filter        L 2230

radiometer with 120 cm of collecting area and have found it

would still require a field size of at least 6 milliradians to

insure adequate NETD (noise equivalent temperature difference).

We are providing two sets of procurement specifications.  The

first procurement specification utilizes an uncooled selected

D* PYROELECTRIC detector and our standard 6-inch SA-1000

radiometer system.  The second procurement specification

utilizes a liquid-nitrogen-cooled MCT detector with our

standard 4-inch SA-1000 radiometer system.

L 2240

***********************************************************************

Specification for PYROELECTRIC detector and 6 inch optic system:

***********************************************************************

a) The radiometer system shall utilize a 6-inch diameter

Cassegrain mirror with a 120 cm collecting area.

b) All optics utilized in the system shall be mirrors, where

possible, to insure lowest system losses.

L 2250

c) The system shall utilize a temperature-stabilized (45�C)

PYROELECTRIC detector of selected D* 3.0e8.

d) A heated BB reference (45�C) shall be used to provide a

-- 30.0 -- In-depth Design of a Filter Radiometer for Cloud Viewi... ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   50 --

stable reference and, through proper phase switching, allow

heated blackbodies to be utilized for calibration testing.

e) The field of view shall be of a precision mask (user

changeable).  Two masks shall be supplied:  2.5 mr

(milliradian), and 6.0 mr.                                       L 2260

f) The sights shall consist of one rifle scope for quick

alignment, and one manually operated through the optics sight.

g) The instrument's field of view shall be described by the

limits of the field where the signal level falls to 50% of the

peak signal (mid-field) when scanned by a modulated source.

The instrument's chopper is stopped in the open position during

the field profile scanning.  The aperture size used on the

source shall be a compromise between field resolution and S/N.   L 2270

The source aperture shall be as small as possible in

combination with source temperature to achieve a mid-field peak

signal level of minimum 200/1.

h) The instrument's actual FOV shall be desired value +/- 10%.

The instrument's field flatness shall be at least 90% of the

peak response over 80% of the FOV.

i) The electrical bandwidth shall be variable in two steps ,

1Hz, and 3 Hz.                                                   L 2280

j) The electronics and data processing unit shall be an IBM

type compatible computer, with 486 cpu, 8 megs of onboard

memory, a hard disk with a capacity of 250 megabytes and a 1.44

meg 3.5-inch floppy disk, keyboard, and mouse.  The monitor

shall be a VGA 14-inch color monitor.

k) The electronics shall use RS-232 or RS-422 serial link for

control and data communications between radiometer head and

electronics unit.                                                L 2290

l) The electronics shall be capable of displaying the

radiometer data in �C over a range from -90 to +10 �C.  The

electronics shall be capable of logging temperature data to the

hard disk at user-selected time intervals.

m) The complete radiometer system shall have an NETD of 1 �C or

less in the 3 Hz bandwidth.  Stability of the radiometer over

the environmental temperature range shall be less than +/- 0.1

�C.  The system shall reach maximum stability within 30 minutes  L 2300

-- 30.0 -- In-depth Design of a Filter Radiometer for Cloud Viewi... ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   51 --

after startup.

n) The interconnecting cable between radiometer head and

electronics shall be 15 feet in length.

o) The instruction manual shall contain all appropriate

operation, set-up, and calibration procedures, as well as block

diagrams and schematics.  Care and cleaning procedure for the

mirror optics shall be included.  The FOV plots vertical and

horizontal, as well the measured net in each electrical          L 2310

bandwidth.

p) The system shall be capable of proper operation while the

power line voltage is within 110-125 volts, frequency 58-62 Hz.

q) The computer shall run under a Windows environment using

Visual Basic and/or Visual C++ as the operating program.

r) The system shall be delivered with a warranty agreement

covering parts and labor for a period of 12 months.  The         L 2320

warranty period shall start at the time of delivery, providing

the system meets the specifications.  The buyer shall have two

weeks to test and evaluate system performance.  If, during the

warranty period, the system fails or ceases to meet

specifications, the user shall notify Minarad Spectrum by FAX

or writing, effectively stopping the warranty clock.  Once the

unit is returned to the buyer and buyer rechecks the system

operation, they shall notify Minarad by FAX or mail, thus

restarting the warranty clock.

L 2330

Minarad is not responsible for buyer's gross abuse, negligence,

and/or mishandling.  The warranty shall cover normal wear and

tear, and shipping and handling.

**************** PYRO detector Sensitivity Worksheet********************

The following calculations were used to estimate the S/N

(signal-to-noise ratio) of this system.  In general, only

theoretical designs with S/N's greater than 100 have been shown

to result in functional systems.                                 L 2340

System:  MSI SA-1000 Filter Radiometer with PYRO

Temp(�C):  -90

Filter Wavelength (microns):  10.8 filter

-- 30.0 -- In-depth Design of a Filter Radiometer for Cloud Viewi... ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   52 --

Spectral Bandwidth (microns):  1.5

FOV (milliradians):  6.0                                         L 2350

Detector Area (sq mm):  1

Electrical Bandwidth (Hz):  3

Optics:  6 inch Cassegrain

Net Collecting Area (cm):  120

Total Optical efficiency:  0.14                                  L 2360

***** derived values *****

Peak Wavelength (microns):  15.8

Solid angle W_R steradians:  2.8e-5

NEP noise equivalent power in Watts:  5.77e-10

NEPD noise equivalent power density in Watts/cm�:  3.44e-11      L 2370

NEN noise equivalent radiance in Watts/cm�-ster:  1.22e-6

N_EFF effective irradiance in Watts/cm�-ster:  8.44e-5

WET equivalent power at detector in Watts:  4.00e-8

S/N signal-to-noise ratio:  69

NETD noise equivalent temperature difference �C: 0.314           L 2380

Note:  the S/N is 69 and less than 100.  This theoretical

design is likely to be realized in our SA-1000 Filter

Radiometer but performance may be marginal.  The equations and

calculations used to derive these values can be reviewed by

recalling the following template.

#                  � ** If highlighted, type Q to get equations      **

#    IRRACLPY.EQU  � ** discussed above.  Scroll down to highlight.  **   L 2390

-- 30.0 -- In-depth Design of a Filter Radiometer for Cloud Viewi... ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   53 --

***********************************************************************

Informal Specifications for a liquid-nitrogen-cooled MCT 4-inch

Cassegrain-mirror radiometer system:

***********************************************************************

Note:  omitted letters imply specifications identical to those

of the PYROELECTRIC system described above.

L 2400

a) The radiometer system shall utilize a 4-inch diameter

Cassegrain mirror with a 45 cm collecting area.

c) The system shall utilize a liquid-nitrogen-cooled MCT

detector of a minimum D* of 2.0e10.  The detector shall have a

minimum hold-time of 8 hours after initial fill and top off.

e) The field of view shall be of a precision mask (user

changeable).  Two masks shall be supplied:  2.5 mr and 5.0 mr.

L 2410

i) The electrical bandwidth shall be variable in three steps

1Hz, 3Hz, 10Hz.

l) The electronics shall be capable of displaying the

radiometer data in �C over a range from -90 to +10 C.  The

electronics shall be capable of logging temperature data to the

hard disk at user-selected time intervals.

m) The complete radiometer system shall have an NET of .08 � C

or less in the 3 Hz bandwidth (2.5mr FOV).  Stability of the     L 2420

radiometer over the environmental temperature range shall be

less than +/-0.1 � C.  The system shall reach maximum stability

within 30 minutes after start-up.

******************* MCT detector Sensitivity Worksheet******************

The following calculations were used to estimate the S/N

(signal-to-noise ratio) of this system.  In general, only

theoretical designs with S/N's greater than 100 have been shown

to result in functional systems.                                 L 2430

System:  MSI SA-1000 Filter Radiometer with MCT (cooled at 77K)

Temp(�C):  -90

Filter Wavelength (microns):  10.8 filter

Spectral Bandwidth (microns):  1.5

-- 30.0 -- In-depth Design of a Filter Radiometer for Cloud Viewi... ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   54 --

FOV (milliradians):  2.5                                           L 2440

Detector Area (sq mm):  1

Electrical Bandwidth (Hz):  3

Optics:  4 inch Cassegrain

Net Collecting Area (cm):  45

Total Optical efficiency:  0.14                                    L 2450

***** derived values *****

Peak Wavelength (microns):  15.8

Solid angle W_R steradians:  4.9e-6

NEP noise equivalent power in Watts:  8.66e-12

NEPD noise equivalent power density in Watts/cm�:  1.37e-12        L 2460

NEN noise equivalent radiance in Watts/cm�-ster:  2.81e-7

N_EFF effective irradiance in Watts/cm�-ster:  8.44e-5

WET equivalent power at detector in Watts:  2.60e-9

S/N signal-to-noise ratio:  300

NETD noise equivalent temperature difference �C:  7.29e-2          L 2470

Note:  the S/N is greater than 100.  This theoretical design is

likely to be successfully realized in our SA-1000 Filter

radiometer.  The equations and calculations used to derive

these values can be reviewed by recalling the following

template.

#                  � ** If highlighted, type Q to get equations      **

#    IRRACLMC.EQU  � ** discussed above.  Scroll down to highlight.  **   L 2480

***********************************************************************

Additional Comments and explanations:

-- 30.0 -- In-depth Design of a Filter Radiometer for Cloud Viewi... ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   55 --

***********************************************************************

1) A precision in-line heated reference source provides a great

many improvements in system performance, such as (a) absolute

reference value vs.  target radiance, (b) the increased S/N

when sky temperatures are close to the ambient temperature, and  L 2490

(c) the ability as in d) above to switch synchronization

polarity and use absolute values of bb references of hotter

than internal reference temperature for calibration or

calibration check.

2) The software we provide includes Visual Basic source code,

if applicable, running in a Windows environment.  The source

code is licensed by Minarad to the model number and serial

number of the instrument.  Having access to the source code

allows the user to modify the Minarad programs as they wish.     L 2500

3) A table of hotter than reference black body temperature vs.

equivalent cold temperature will be provided to you with the

instrument.

4) Although Minarad has chosen a temperature range we feel will

work for you, please feel free to suggest the temperature range

or ranges you desire, as well as alternate FOV precision field

stop masks.

L 2510

We look forward to your comments on our specifications.

��������������͸

� FIBER OPTICS ���������������������������������������������������������

31.0       Calculation of Decibel Loss and %Transmittance of Fibers

������������������������������������������������������������������������

The MSI VIR-20 spectrometer/spectroradiometer uses fiber light

guides to couple energy from sensors to its internal             L 2520

monochromators.  The VIR-20 has two independent channels:  a

sample or main channel and a reference or auxiliary channel.

The VIR-20's spectral range covers 300 to 2500 nm.  For this

range UV-enhanced water-free quartz fibers are chosen.  For

longer wavelengths, HMFG (heavy metal fluoride glass) and

chalcogenide fibers are appropriate.  The calculations

presented below are useful for estimating the transmittance

factor for a short fiber length.  The transmittance factor is

then included in the overall system throughput factor used to

estimate S/N (signal-to-noise).                                  L 2530

-- 31.0 -- Calculation of Decibel Loss and %Transmittance of Fibe... ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   56 --

Topic # 55: FIBER OPTICS, LOSS DB/KM, %TRANSMITTANCE/METER

Problem # 55.1:

(a) Calculate a fiber's loss in decibels per kilometer LOSS_DB_KM

given a percent transmittance per meter T_%_MTR of 97 and

(b) the percent transmission per meter if the LOSS_DB_KM is 100.

(c) What are the DB loss and % transmittance for a 4 meter length?

#                  � ** If highlighted, type Q to get equations that **   L 2540

#    IRRAFIBL      � ** solve this problem. Scroll down to highlight.**

������������������������͸

� SERIES and INTEGRATION �����������������������������������������������

32.0       Circumference of an Ellipse: Calculation by Series and

Integration

������������������������������������������������������������������������

L 2550

Calculations involving integration and summation of formulas

are relevant to all math-based theories.  The following

examples illustrate these methods.

The circumference of an ellipse with axes lengths of "a" and

"b" as measured from its origin cannot be expressed in a simple

"closed" form by a formula.  The approximate value is

calculated as:

l � 2 * pi * sqr( (a^2 + b^2)/2 )                         L 2560

Thus if a = 50 and b = 10, the circumference, l � 226.5.  The

circumference can be express "exactly" by use of the elliptic

integral of the second kind E(k,�):

l = 4 * a * E(k,�)   k= sqr( abs(a^2-b^2) )/a

There are at least two methods to obtain a higher degree of

accuracy:  the first is to use a series expansion for E(k,�)

and the second is to numerically integrate E(k,�).  The          L 2570

following two examples show how AstroFormCalc 3 can be used to

sum up a series or to accurately integrate an equation.  The

methods are very general and are applicable to many problems in

Optics as well as virtually all other disciplines.  It should

be noted that the "approximation" formula, given above, is only

accurate to about 8%.

----- 32.0 -- Circumference of an Ellipse: Calculation by Series and ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   57 --

Topic # 56: SERIES, ARC LENGTH (CIRCUMFERENCE) OF ELLIPSE

Discussion: Estimate the circumference of an ellipse with axes lengths    L 2580

of 50 and 10 cm as measured from its origin. Use the series

approximation.

#                  � ** If highlighted, type Q to get equations that **

#    IRRAELSE      � ** solve this problem. Scroll down to highlight.**

Topic # 57: INTEGRATION, ARC LENGTH (CIRCUMFERENCE) OF ELLIPSE

Discussion: Estimate the circumference of an ellipse with axes lengths    L 2590

of 50 and 10 cm as measured from its origin. Use the elliptic integral

method.

#                  � ** If highlighted, type Q to get equations that **

#    IRRAELIN      � ** solve this problem. Scroll down to highlight.**

�����������͸

� GLOSSARY  ������������������������������������������������������������

L 2600

33.0     Glossary of Terms

������������������������������������������������������������������������

� Note, Use the Jump command to locate a glossary reference.    �

� Scroll the screen until the reference is in the middle at line�

� 12.  The Jump box will light up.  Type J to locate.  Repeat   �

� typing A, as needed, to continue the search.                  �

Cartesian ovoid.  See  IRRA402 .                                          L 2610

Compound thick lenses.  See  IRRA443 .

Compound thin lenses.  See  IRRA427 .

Integration.  See  IRRAELIN .

IR Source (blackbody).  See  IRRA1946 .

Lens Imagery.  See  IRRA419 .                                             L 2620

Mirrors.  See  IRRA495 .

------------------------------------------ 33.0 -- Glossary of Terms ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   58 --

Series.  See  IRRAELSE .

Spherical refracting surfaces.  See  IRRA405 .

Radiation.  See  IRRASLWV .

Thick Lenses.  See  IRRA442 .                                             L 2630

Thin Lenses.  See  IRRA410 .

�����������������͸

� WANT TO SUBMIT? ������������������������������������������������������

34.0       MSI Encourages Users to Submit Calculations.

������������������������������������������������������������������������

The on-line manual studies many scientific concepts which when   L 2640

grouped together are commonly called "Optics and IR

Radiometry".

This manual is designed for worldwide distribution over the

"digital superhighway" as shareware.  It is very portable and

will run on virtually any PC.  It will be uploaded to many BBS

and on-line services such as America on-line, CompuServe,

Prodigy, ...  It will also be included on the ASP CD-ROM and

many other engineering and scientific CD-ROMS.

L 2650

MSI also welcomes users of this file to submit their own

examples for possible use in future releases.

All submissions should include equations written in the style

of the *.EQU files included with this manual.  The many *.EQU

files can serve as templates.  Also, refer to the

AstroFormulaCalculator 3 on-line helpfile (F1 key) for detailed

instructions on entering equations and making calculations.

Submissions should also include a comprehensive, but concise,

text for inclusion in this manual.  MSI will give full author    L 2660

credit to all submissions which are accepted for this manual.

Please submit these items on floppy media to the Minarad

address in the "About Minarad " section, below.

���������͸

� LICENSE ��������������������������������������������������������������

35.0         Product License and Warranty.

------------------------------ 35.0 -- Product License and Warranty. ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   59 --

������������������������������������������������������������������������

L 2670

Software License Agreement  (Registered Users)

If you are a registered user, Minarad Spectrum, Inc.  (MSI)

grants you the right to use this Software for a single-user

computer or a workstation not attached to a network.  You may

not rent or lease this Software without the expressed written

permission of MSI.  You may not decompile, disassemble, reverse

engineer, copy, create a derivative work, or otherwise use this

Software except as stated in this agreement.  The license

authorizes you to create and use an extra copy of the Software   L 2680

on a home or laptop computer, as long as the extra copy is

never loaded at the same time this Software is loaded on the

primary computer on which you use this Software.

If any provision of this agreement is found to be unlawful,

void, or unenforceable, then that provision shall be severed

from this Agreement and will not affect the validity and

enforceability of any of the remaining provisions.  This

agreement shall be governed by the laws of the State of

Delaware.                                                        L 2690

The Software and accompanying materials are provided with

Restricted Rights.  Use, duplication, or disclosure by the

Government is subject to restrictions as set forth in

subparagraph (c) (1) (ii) of The Rights in Technical Data and

Computer Software clause at DFARS 252.227-7013, or

subparagraphs (c) (1) and (2) of the Commercial Computer

Software--Restricted Rights at 48 CFR 52.227-19, as applicable.

Contractor/manufacturer is Minarad Spectrum, Inc., at the

address listed below.                                            L 2700

Non registered users must comply with the Shareware policy

described elsewhere in this software Package.

Type T to return to the Table of Contents.

Warranty Information

Minarad Spectrum, Inc.  (MSI) hereby disclaims all warranties

relating to this Software, whether express or implied,           L 2710

including without limitation any implied warranties of

merchantability or fitness for a particular purpose.  MSI will

not be liable for any special, incidental, consequential,

indirect or similar damages due to loss of data, loss of

------------------------------ 35.0 -- Product License and Warranty. ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   60 --

profits, lost savings or any other reason, even if MSI or an

agent of MSI has been advised of the possibility of such

damages.  In no event shall MSI's liability for any damages

ever exceed the price paid for the license to use Software,

regardless of the form of the claim.  The person using the

Software bears all risk as to the quality and performance of     L 2720

the Software.  This Software is licensed AS IS.

Any action or proceeding brought by either party against the

other arising out of or related to this agreement shall be

brought in a STATE or FEDERAL COURT of competent jurisdiction

located in Fairfield County, Connecticut.  The parties hereby

consent to in personam jurisdiction of said courts.

Type T to return to the Table of Contents.

L 2730

��������������͸

� USER SUPPORT ���������������������������������������������������������

36.0                    User Support.

������������������������������������������������������������������������

This program has undergone over 5 years of "beta testing"

around the world.  The algorithms have been in use since the

1960's, are proven and believed highly reliable.  However,

comprehensive programs of this nature may develop what are       L 2740

called "bugs".  User support is provided to all registered

users for a minimum of 3 months without charge.  Registered

users can receive such help via phone/mail.  Registration also

enrolls the user for notification of "bug" fixes, upgrades, new

versions, new products and many more benefits.  The phone and

fax numbers for Minarad Spectrum are:

Voice: 203 371-6348        Fax: 203 371-6349

To jump to the registration section, position the jump line in   L 2750

the center of the display and type J.   �Registration� .

Support will be provided to non-registered users and users

beyond the free support period on a time and materials basis.

��������������͸

� REGISTRATION ��������������������������������������������������������

37.0                 Order/Registration Form

������������������������������������������������������������������������  L 2760

------------------------------------ 37.0 -- Order/Registration Form ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   61 --

If a printer is available, position the order form (part 1) on

the screen and type the "(print screen)" key to make a copy.

Then, advance the text and position the order form (part 2) on

the screen and type "(print screen)" again to finish.

Minarad Spectrum Products Order/Registration Form  (part 1)

Name:________________________________  Title:________________

L 2770

Company:_____________________________________________________

Address:____________________________ City:___________________

State:___________   Zip:_____________  Country:______________

Date:_____/_____/_____           (Fill out form completely)

---------------------------------------------------------------

Product                     #     Price/ea     S&H       Total        L 2780

---------------------------------------------------------------

OPTICS, IR and RADIOMETRY   ___    _$29.95_    _$5.00_   _______

Calculations By Nelson

Engborg and Paul Murphy

------------------------------------------------------------------------

------------------------------------------------------------------------  L 2790

Minarad Spectrum Order/Registration Form  (part 2)

* Canadian Orders, Add $2.00 per copy                   _______

* International Orders, Add $6.00 per copy              _______

* CT Residents add appropriate sales tax                _______

TOTAL ENCLOSED      __________

L 2800

Disk Size: [ ] 5-1/4"   [ ] 3-1/2"   [ ] Check  [ ] Visa  [ ] MC

Card# __________________________________     Exp. Date ________

Signature ______________________________

------------------------------------ 37.0 -- Order/Registration Form ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   62 --

How did you hear about our Products?    [ ] BBS   [ ] Friend

[ ] Other _____________________________________________________

Mail Completed form with Payment to:                    L 2810

Minarad Spectrum, Inc.   Dept. MSI312

7 Cambridge Drive.

Trumbull,  CT   06611    USA

Support ONLY,  203 371-6348     Fax ONLY, 203 371-6349

[Aug 1995]  Prices subject to change without notice.

Type T to return to the Table of Contents.                 L 2820

������������͸

� REFERENCES �����������������������������������������������������������

38.0           Optics and Infrared References.

������������������������������������������������������������������������

M.  Born and E.  Wolf, "Principles of Optics", Pergamon, Oxford, 1980.

Dr.  M.  Foigel, Chief Editor, "The Optics Problem Solver", Research     L 2830

and Education Association, 61 Ethel Road West, Piscataway, New Jersey,

08864, Revised Printing 1990

E.  Hecht, "Optics", Addison-Wesley, Reading, Massachusetts, 1988.

Eugene Hecht, Ph.D., "Optics", Schaum Outline of Theory and Problems,

McGraw-Hill Book Company, New York, 1975

M.  V.  Klein and T.  E.  Furtak, "Optics", John Wiley & Sons, New

York, 1986.                                                              L 2840

R.  S.  Longhurst, "Geometrical and Physical Optics", John Wiley &

Sons, New York, 1964.

F.  W.  Sears, " Optics", Addison-Wesley, Cambridge, Massachusetts,

1948.

------------- IR References -------------

Nelson Engborg, Al Kozlowski and Paul Murphy, "IR polarization:          L 2850

Applications to sources and radiometers", Technical Note #11, Minarad

Spectrum, Inc., May 1994.

---------------------------- 38.0 -- Optics and Infrared References. ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   63 --

Nelson Engborg and Paul Murphy, "Discussion of Field of View (FOV),

spot diameter and working distance for an infrared system", Technical

Note #21, Minarad Spectrum, Inc., May 1994.

Nelson Engborg and Paul Murphy, "Non-contact temperature measurement

with the infrared microscope", Technical Note #31, Minarad Spectrum,

Inc., May 1994.                                                          L 2860

Al Kozlowski, Al Waldie, Nelson Engborg and Paul Murphy, "8 inch Target

Projector:  Collimator, target wheel and IR source", Technical Note

#41, Minarad Spectrum, Inc., May 1994.

Gerald C.  Holst, "Testing and Evaluation of Infrared Imaging Systems",

JCD Publishing Co., 2923 Cove Trail, Maitland FL 32751, 1993.

W.  L.  Wolfe and G.  J.  Zissis, "The Infrared Handbook", Infrared

Information and Analysis Center (IRIA), Environmental Research           L 2870

Institute of Michigan, Revised Printing, 1985.

���������������������͸

� MANUALLY INSTALLING ��������������������������������������������������

39.0       Instructions for Manually Installing Software Package.

������������������������������������������������������������������������

The following are minimal instructions for manually installing

Minarad Spectrum's IR Radiometry and Optics Package.                 L 2880

Minimum distribution files:  README.1ST, MSI312.ICO, ACAL3.EXE,

ACAL3HLP.HLP and MSI312.HLP with its *.EQU and other files.

These files will probably be bundled in a compressed file such

as MSI312.ZIP, MSI312.SDN, or something similar.

Create a DOS directory named MSI and change to it:  cd\ (enter)

md\msi (enter) cd \msi (enter)

Please unload.  For example, PKUNZIP MSI312.ZIP (enter).         L 2890

ACAL3.EXE, ACAL3HLP.HLP, MSI.ICO and MSI312.HLP and all *.EQU

files should reside into the directory MSI.  Copy MSI312.BAT

into root.

Add these statements to AUTOEXEC.BAT in the base directory:

SET ASTRO3=C:\MSI. Append to existing PATH statement ;C:\MSI;

Running Under Windows as a Non-Windows Application

----- 39.0 -- Instructions for Manually Installing Software Package. ---

---------- MSI312.HLP ----------- Minarad Spectrum -------------   64 --

Follow the installation instructions and install on the hard     L 2900

drive C.  Start WINDOWS (tm, MicroSoft).  Open the Accessory

Window.  Click on FILE then click on PROPERTIES.  Enter the

following information for each of the fields.  Description:  IR

Radiometry, Command:  ACAL3.EXE MSI312.HLP, Working Directory:

\MSI.  Click on CHANGE ICON.  Enter Filename:

C:\MSI\MSI312.ICO.  Click on OK to exit and click on OK to exit

FILE.  The MSI icon should appear in the accessory window.

Click on it to run MSI312.

To start from DOS:  cd\msi (enter)    acal3 MSI312.hlp (enter)     L 2910

Access on-line help (file: ACAL3HLP.HLP) from equations

display with (F1) key.

�����������͸

� HELP FILE ������������������������������������������������������������

40.0          Help File: Summary of File Viewing Commands

������������������������������������������������������������������������

L 2920

A - Auto-search of entire text for user-supplied string. Search is

in uppercase and is case insensitive. If string found,

type A to continue or to enter new string, type (space) A.

B,N - Move back or ahead 18 lines.  Also use (backspace) for "B".

E, (esc) - Exit from help file to equations display.

F,T,K - Jump to the first (beginning), T.O.C or last (end) line.

G - Skip to Glossary of terms or summary of functions, if present.

J,(ent) - Search line at <JUMP�, Reference is after  . Jump to

line number or perform autosearch ahead if text reference.

L,P - Jump to line/page. Type Lxxx(ent) or Pyy(ent) to jump to        L 2930

line #xxx or page #yy. With + or -, ahead or back 56 lines.

S,R - Skip ahead (S) or back (R) in help file to major headings

W - Skip to Summary of useful calculations or warnings, if present

1 to 9 - Typing 1 to 9 advances the text by 1 to 9 lines.

(ctrl Q) - Exit from help file.  Same as command "E".

Edit keys - Use edit keys: (Home) (up arrow) (PgUp) (left arrow)

(right arrow) (End) (down arrow) (PgDn).

Entry into the application file is at/near the last position.

Type: T to return to Table of Contents or E to exit the file.        L 2940

---------------- 40.0 -- Help File: Summary of File Viewing Commands ---